Supplier name: Micron Equip

Supplier contact: andile@micronequip.com

Product/service supplied:

The laboratory has designated Micron Equip as the exclusive vendor for the procurement, servicing, and calibration of the equipment outlined in the Table below:

| Equipment | Description | Calibration Required? |
| --- | --- | --- |
| 25mm Crusher | 25mm 3-phase jaw crusher, capable of crushing ±50mm coal down to size -25mm. The crusher must have adjustable jaws, have a guard mesh housing all mechanical components, and a ±100mm receiving mouth. | No, efficiency to be conducted by the lab |
| 10mm Crusher | 10mm 3-phase jaw crusher, capable of crushing ±25mm coal down to size ±25mm. The crusher must have adjustable jaws, have a guard mesh housing all mechanical components, and a ±50mm receiving mouth | No, efficiency to be conducted by the lab |
| 3mm Crusher | 3mm 3-phase jaw crusher, capable of crushing ±10mm coal down to size ±3mm. The crusher must have adjustable jaws, have a guard mesh housing all mechanical components, and a ±20mm receiving mouth. | No, efficiency to be conducted by the lab |
| Pulverisers | Pulverisers capable of milling coal from -3mm to -212µm. They must either use a cross-beat or hammers, possess a 212µm sieve with a receiver. Both 3-phase & 2-phase are acceptable. | No, efficiency to be conducted by the lab. |
| Homogenizer | A 3-phase homogenizer capable of processing a 210 L drum of pulverised coal. | No, efficiency done by the lab. |
| Oven | A 700L-1400L oven capable of reaching temperatures of 110-120 ºC with ±2 heating fans. | Yes, at setpoints 40ºC & 107ºC |
| Oven | A 50L oven capable of reaching temperatures of ±110ºC. | Yes, at setpoint 107ºC |
| Furnace | A muffle furnace capable of reaching temperatures of up to ±815ºC | Yes, at setpoints 500ºC & 815ºC, ramping profile. |
| Furnace | A muffle furnace capable of reaching temperatures of up to ±900ºC. | Yes, at setpoint 900ºC |
| Sieve Screens | - Sieve screens with a 450mm circumference and sizes in mm (50,40,25,22.4,10,6,4.75,3,1,0.5). - Sieve screen with ±100mm circumference and size 212µm. - 450mm circumference receiver. - ±100mm circumference receiver. | No, verification done by the lab. |
| Drycal Calorimeter System | - Drycal Single Vessel Calorimeter System - Single Vessel Drycal Calorimeter Complete with: 1 x Computer complete 1 x Calorimeter 1x Controller box 2x Bombs 1x Filling station 1x Cooler/ washer 1x Circulation Tank 1x UPS 1x Startup seal kit 1x Operation manual - Water Chiller System Complete - Drycal Oxygen Regulator Complete Fittings & Piping - Drycal Crucibles - Benzoic Acid Tablets, mass at ±0.5g, CV at 26.423 MJ/KG | No, calibration to be done by the lab using Benzoic acid tablets. |
| Analytical Balance | Capacity: 120g Readability: 0.1mg Pan Size: 80mm x 80mm | Yes |
| Top Pan Scale | Capacity: 8000g Readability: 0.1g Pan Size: 300mm x 210mm | Yes |
| Platform Scale | Capacity: 150kg Readability: 10g Pan Size: 400mm x 500mm | Yes |
| Mass Pieces | Class F1 Weights(100g,50g,20g,10g,5g,1g,0.5g) Mass Piece Cast Iron (20kg, 10kg, 5kg) | Yes |
| Cone & Quarter Cross | A cone & quarter steel cross at ±7m for both the legs | No |
| Desiccator | ±20L desiccator, with moisture-absorbing silica gel desiccants. | No |

Note: Consumables for any equipment purchased at Micron shall also be purchased at Micron Equip.

To be filled by the service provider

I, ___________________________________________ on behalf of Micron Equip hereby mandate that Micron Equip will be able to undertake the Supply, Maintenance & Calibration of the aforementioned equipment.

Signature: ________________________

Date: ____________________________

To Be Filled By Azuresky Pty Ltd

Competence & Acceptance Criteria

The overall competence & acceptance criteria required for the purchase & calibration of equipment is a SANAS accreditation (mass & temperature metrology) & ± 5 years of supply, servicing & calibration of equipment.

Remarks: This appointment is based on Micron Equips’ demonstrated technical competence and a proven track record of reliability established through previous institutional engagements. Consequently, Micron Equip is recognised as a SANAS-accredited, pre-approved, trusted supplier for all laboratory technical requirements.

Supplier approval

☒ Approved
☐Conditionally Approved (e.g., for low-risk items or pending documents)
☐Not Approved

Evaluator Name: Tshepo Maponya

Signature: ___________________________ 
Date: ________________________________