## PRECISION
|  | Analyst 1 (Average) 04/07 |  | Analyst 2 (avergae)
Set | LIMS NO | Mass: 0.5000-0.5099 (g) | LIMS NO | Mass: 0.5000-0.5099 (g)
1 | 1168 | 23.814 | 1303 | 23.795
2 | 1169 | 23.883 | 1304 | 23.76
3 | 1170 | 23.9695 | 1305 | 23.71
4 | 1171 | 23.974 | 1306 | 23.825
5 | 1172 | 23.871 | 1307 | 23.735
6 | 1173 | 23.949 | 1308 | 23.815
7 | 1174 | 23.8735 | 1309 | 23.71
8 | 1175 | 24.118 | 1310 | 23.82
Average |  | 23.9315 |  | 23.7712
Absolute difference |  | 0.16025
Std deviation |  | 0.093749 |  | 0.048825
RSD |  | 0.39174 |  | 0.2054
Assigned Value (AD base) |  | 24.07
The RSD values of the masurements indicate that there is low variability and high precision within the data set. This is a good indication that the data is clustered relatively closer to the mean
Precision is confirmed.

## ACCURACY
Using the data under the PRECISION SHEET
| Analyst 1 | Analyst 2
Average | 23.9315 | 23.7712
Assigned value | 24.07
Absolute difference | 0.1385 | 0.29875
Acceptance criteria ≤0.30 MJ/KG
Recovery | 99.4246 | 98.7588
Since the absolute differences do not exceed the 0.3 MJ/kg limit and the recoveries are so high, the method is deemed accurate.

## REPEATABILITY AND REPRODUCIBILI
|  | Analyst 1 (Average) 04/07 |  | Analyst 2 (average) | LIMS NO | Analyst 3 (Average)
Set | LIMS NO | Mass: 0.5000-0.5099 (g) | LIMS NO | Mass: 0.5000-0.5099 (g) |  | Mass: 0.5000-0.5099 (g)
1 | 1168 | 23.814 | 1303 | 23.795 | 1366 | 23.85
2 | 1169 | 23.883 | 1304 | 23.76 | 1367 | 23.89
3 | 1170 | 23.9695 | 1305 | 23.71 | 1368 | 23.92
4 | 1171 | 23.974 | 1306 | 23.825 | 1369 | 23.92
5 | 1172 | 23.871 | 1307 | 23.735 | 1370 | 23.95
6 | 1173 | 23.949 | 1308 | 23.815 | 1371 | 24.01
7 | 1174 | 23.8735 | 1309 | 23.71 | 1372 | 23.94
8 | 1175 | 24.118 | 1310 | 23.82 | 1373 | 23.95
Average |  | 23.9315 |  | 23.7712 |  | 23.9287
Std Dev |  | 0.093749 |  | 0.048825 |  | 0.047037
Assigned Value |  | 24.07
Abs Difference |  | 0.1385 |  | 0.29875 |  | 0.14125
Acceptance criteria: | Difference b/w results and assigned value should be ≤0.30 MJ/kg
Repeatability and reproducibility conditions are met

## LINEARITY
Since we're comparing two conditions (~0.3 and ~0.5 g) with the same sample set, linearity isn't directly applicable unless these are collected over a concentration or time gradient. In this case, the lab is validating method robustness/ruggedness.

## RUGGEDNESS-ROBUSTNESS
2025-07-04
LIMS no | Normal Conditions (19-26 oC) | Elevated temperature (average temp of 27.56oC) (08-09 July 2025) | LIMS no | Anova: Single Factor
1168 | 23.814 | 23.8315 | 1356
1169 | 23.883 | 24.0245 | 1357 | SUMMARY
1170 | 23.9695 | 23.8705 | 1358 | Groups | Count | Sum | Average | Variance
1171 | 23.974 | 23.7995 | 1359 | Normal Conditions (19-26 oC) | 8 | 191.452 | 23.9315 | 0.0087888
1172 | 23.871 | 23.939 | 1360 | Elevated temperature (average temp of 27.56oC) (08-09 July 2025) | 8 | 191.013 | 23.8766 | 0.0052021
1173 | 23.949 | 23.8575 | 1361
1174 | 23.8735 | 23.8575 | 1362
1175 | 24.118 | 23.833 | 1363 | ANOVA
|  |  |  | Source of Variation | SS | df | MS | F | P-value | F crit
F-crtical | 4.6001 |  |  | Between Groups | 0.012045 | 1 | 0.012045 | 1.7218 | 0.21056 | 4.6001
F-test | 1.7218 |  |  | Within Groups | 0.097936 | 14 | 0.0069954
p-value | 0.21056
|  |  |  | Total | 0.10998 | 15
Since F-stat is less than the F-critical, and also the p-value is >0.05. There is no significant difference between the different conditions
Which implies that the method passes the ruggedness/robustness assessment

## HOMOGENEITY ASSESSMENT
RM name | GA CRM AS002
Measurement Method | ISO 1928
Date of characterisation | 2025-05-01
Analyst | Keitumetse
Number of replicates | 10 units * 2 replicates
Unit | Replicate 1 | Replicate 2 | Unit Mean | Between SS | Within SS
Unit 1 | 23.814 | 23.795 | 23.8045 | 0.0087384 | 0.0001805 | Anova: Single Factor
Unit 2 | 23.883 | 23.76 | 23.8215 | 0.0048216 | 0.0075645
Unit 3 | 23.9695 | 23.71 | 23.8398 | 0.0019034 | 0.03367 | SUMMARY
Unit 4 | 23.974 | 23.825 | 23.8995 | 0.0016704 | 0.011101 | Groups | Count | Sum | Average | Variance
Unit 5 | 23.871 | 23.735 | 23.803 | 0.0091395 | 0.009248 | Replicate 1 | 10 | 239.442 | 23.9442 | 0.0075581
Unit 6 | 23.949 | 23.815 | 23.882 | 0.00025992 | 0.008978 | Replicate 2 | 10 | 237.97 | 23.797 | 0.0056011
Unit 7 | 23.8735 | 23.71 | 23.7918 | 0.012435 | 0.013366
Unit 8 | 24.118 | 23.82 | 23.969 | 0.019365 | 0.044402
Unit 9 | 23.99 | 23.84 | 23.915 | 0.0039427 | 0.01125 | ANOVA
Unit 10 | 24 | 23.96 | 23.98 | 0.023937 | 0.0008 | Source of Variation | SS | df | MS | F | P-value | F crit
[6 empty] | Between Groups | 0.10834 | 1 | 0.10834 | 16.4659 | 0.00073838 | 4.4139
[6 empty] | Within Groups | 0.11843 | 18 | 0.0065796
|  |  | 23.8706 | 0.086213 | 0.14056 | Total | 0.22677 | 19
|  |  | df_B: | 9
|  |  | df_W: | 10
|  |  | MSB: | 0.0095792
|  |  | MSW: | 0.014056
|  |  | s_bb^2: | 0
|  |  | s_bb: | 0
|  |  | Repeatability SD: | 0.11856
|  |  | Fcalc: | 0.6815
|  |  | Fcrit (α=0.05): | 3.0204
|  |  | Conclusion: | If Fcalc > Fcrit → Significant between-unit variation
Since Fstat< Fcritical, Homogeneity Assesment passes

## UoM
1. Uncertainty sources table
Sources of UoM | Symbol | Type (A/B) | Standard Uncertainty (U) | U^2 | Distribution | Divisor | Sensitivity coefficient
Repeatability (within lab) | U1 | A | 0.12 | 0.0144 | Normal | 1 | 1
Balance Calibration | U2 | B | 0.015701 | 0.00024652 | Rectangular | √3
Homogeneity test (b/w units) | U3 | A | 0.10834 | 0.011737 | Normal | 1 | 1
Stability (predicted drift) | U4 | B | NA |  | Rectangular | √3 | 1
Uchar of PT | U5 | B | NA |  | Normal | 1 | 1
2. Combined Uncertainty
To calculate Uom of balance
Combined uncetainty (Uc) | 0.16243 | UoM on calibration certifcate | 0.0003
|  | Average mass of sample | 0.505
|  | Benzoic AV | 26.43
3. Expanded Uncertainty |  | UoM | 0.015701
Expanded uncertainty (k=2) | 0.32486
4. Final Reporting statement
Expanded Uncertainty (k=2) | 0.32486
Confidence level: | 95%
Traceability | All components traceable to SI unit via calibrated instruments