## Sheet1
Title Page
Method Validation Report: Inherent Moisture (SANS 5925)
Laboratory name, date, analysts
1. Introduction
Purpose of validation (to confirm reliability of SANS 5925 for inherent moisture)
Scope (coal samples, inherent moisture, mg range, reference value used = 3.4%)
2. Methodology
Reference method: SANS 5925
Analysts: Keitumetse and Lilo
Dates and run times (20 May, 21 May, 13 June, 17 June 2025)
Replicates: 2 replicates per sample, 10 samples per run
Conditions: Normal laboratory environment
Acceptance criteria (based on ISO 5725 & SANAS TR 81):
Repeatability RSD ≤ 2%
Reproducibility RSD ≤ 5%
Bias ≤ ±0.3% absolute against CRM/reference value
r-limit (repeatability limit) and R-limit (reproducibility limit) respected
3. Results
3.1 Precision (Repeatability)
Within-run %RSD (sr) reported for each analyst/date (0–2.1%) — all within acceptance.
Pooled repeatability standard deviation (sr) = 0.035.
Repeatability limit (r) = 0.098 — observed differences well below limit.
3.2 Reproducibility
Between-run SD of lab means (sR) = 0.214.
Reproducibility limit (R) = 0.598.
All run means fall within R-limit.
3.3 Accuracy and Bias
Reference value: 3.40%.
Mean results:
Keitumetse (20 May): 3.48% (bias +0.08)
Keitumetse (21 May): 3.23% (bias -0.17)
Lilo (13 June): 3.07% (bias -0.33)
Lilo (17 June): 3.08% (bias -0.32)
Overall mean = 3.21% (bias -0.19 vs CRM).
Bias <0.3% abs → within acceptance.
3.4 Robustness
Different analysts, days, and run durations tested → no significant differences that exceed R-limit.
3.5 Measurement Uncertainty
Sources considered: repeatability, reproducibility, balance calibration, environmental effects.
Combined standard uncertainty ≈ 0.11%.
Expanded uncertainty (k=2, 95% CL) ≈ 0.22%.
4. Discussion
Repeatability and reproducibility meet SANAS/ISO 5725 criteria.
Bias is small and within the ±0.3% acceptance.
Method performance is stable across different analysts, days, and conditions.
5. Conclusion
The SANS 5925 method for inherent moisture determination is validated.
It meets SANAS requirements for precision, accuracy, reproducibility, and uncertainty.
Fit for purpose for routine laboratory use.
6. Final Reporting Statement
“The determination of inherent moisture in coal samples by SANS 5925 has an expanded uncertainty of 0.22% (k=2, 95% confidence level). The method is traceable to SI units through calibrated balances and reference materials. The method is fit for SANAS-accredited use.”
Actions: check instrument calibration and maintenance records, review sample preparation SOP, repeat selected runs, and consider applying bias correction or including bias in uncertainty if justified.

## PRECISION-INT PREC & REPR
Date: | 21-05-2025 | Analyst: Keitumetse | Duration: | 1hr |  |  | Date: | 20-05-2025 | Analyst: Keitumetse | Duration: | 1h30 |  |  | Date: | 13-06-2025 | Analyst: Lilo | Duration: | 1h00 |  |  | Date: | 17-06-2025 | Analyst: Lilo | Duration: | 1h30
LIMS no | 1 | 2 | Difference | Keitumetse Average | Standard deviation | %sr (RSD) | LIMS no | 1 | 2 | Difference | Keitumetse Average | Standard deviation | %sr (RSD) | LIMS no | 1 | 2 | Difference | Lilo Average | Standard deviation | %sr (RSD) | LIMS no | 1 | 2 | Difference | Lilo Average | standard deviation | %sr (RSD)
168 | 3.33 | 3.3 | 0.03 | 3.315 | 0.021213 | 0.63992 | 158 | 3.59 | 3.59 | 0 | 3.59 | 0 | 0 | 1168 | 3.11 | 3.02 | 0.09 | 3.065 | 0.06364 | 2.0763 | 1218 | 3.17 | 3.14 | 0.03 | 3.155 | 0.021213 | 0.67237
169 | 3.43 | 3.45 | 0.02 | 3.44 | 0.014142 | 0.41111 | 159 | 3.46 | 3.39 | 0.07 | 3.425 | 0.049497 | 1.4452 | 1169 | 3.06 | 3.06 | 0 | 3.06 | 0 | 0 | 1219 | 2.89 | 2.98 | 0.09 | 2.935 | 0.06364 | 2.1683
170 | 3.19 | 3.23 | 0.04 | 3.21 | 0.028284 | 0.88113 | 160 | 3.47 | 3.48 | 0.01 | 3.475 | 0.0070711 | 0.20348 | 1170 | 3.06 | 3.15 | 0.09 | 3.105 | 0.06364 | 2.0496 | 1220 | 3.12 | 3.12 | 0 | 3.12 | 0 | 0
171 | 3.42 | 3.34 | 0.08 | 3.38 | 0.056569 | 1.6736 | 161 | 3.45 | 3.49 | 0.04 | 3.47 | 0.028284 | 0.81511 | 1171 | 3.14 | 3.11 | 0.03 | 3.125 | 0.021213 | 0.67882 | 1221 | 3.07 | 3.03 | 0.04 | 3.05 | 0.028284 | 0.92735
172 | 3.24 | 3.29 | 0.05 | 3.265 | 0.035355 | 1.0829 | 162 | 3.58 | 3.6 | 0.02 | 3.59 | 0.014142 | 0.39393 | 1172 | 3.13 | 3.11 | 0.02 | 3.12 | 0.014142 | 0.45327 | 1222 | 3.69 | 3.64 | 0.05 | 3.665 | 0.035355 | 0.96468
173 | 3.23 | 3.19 | 0.04 | 3.21 | 0.028284 | 0.88113 | 163 | 3.48 | 3.45 | 0.03 | 3.465 | 0.021213 | 0.61221 | 1173 | 3.09 | 3.07 | 0.02 | 3.08 | 0.014142 | 0.45916 | 1223 | 3.03 | 2.97 | 0.06 | 3 | 0.042426 | 1.4142
174 | 3.17 | 3.22 | 0.05 | 3.195 | 0.035355 | 1.1066 | 164 | 3.53 | 3.63 | 0.1 | 3.58 | 0.070711 | 1.9752 | 1174 | 3.02 | 2.99 | 0.03 | 3.005 | 0.021213 | 0.70593 | 1224 | 3 | 3.03 | 0.03 | 3.015 | 0.021213 | 0.70359
175 | 3.2 | 3.25 | 0.05 | 3.225 | 0.035355 | 1.0963 | 165 | 3.46 | 3.44 | 0.02 | 3.45 | 0.014142 | 0.40992 | 1175 | 2.98 | 2.97 | 0.01 | 2.975 | 0.0070711 | 0.23768 | 1225 | 2.89 | 2.96 | 0.07 | 2.925 | 0.049497 | 1.6922
176 | 3 | 2.99 | 0.01 | 2.995 | 0.0070711 | 0.2361 | 166 | 3.39 | 3.37 | 0.02 | 3.38 | 0.014142 | 0.41841 | 1176 | 3.05 | 3.05 | 0 | 3.05 | 0 | 0 | 1226 | 2.93 | 3.02 | 0.09 | 2.975 | 0.06364 | 2.1391
177 | 3.12 | 3.09 | 0.03 | 3.105 | 0.021213 | 0.68319 | 167 | 3.36 | 3.4 | 0.04 | 3.38 | 0.028284 | 0.83681 | 1177 | 3.12 | 3.05 | 0.07 | 3.085 | 0.049497 | 1.6045 | 1227 | 2.9 | 2.97 | 0.07 | 2.935 | 0.049497 | 1.6865

## BIAS
Run | LIMS | R1 | R2 | Difference | Mean | Within SD | %RSD | Bias vs CRM | Reference Value | 3.4
Keitumetse 21-05 (1 hr) | 168 | 3.33 | 3.3 | 0.03 | 3.315 | 0.021213 | 0.63992 | -0.085
Keitumetse 21-05 (1 hr) | 169 | 3.43 | 3.45 | 0.02 | 3.44 | 0.014142 | 0.41111 | 0.04
Keitumetse 21-05 (1 hr) | 170 | 3.19 | 3.23 | 0.04 | 3.21 | 0.028284 | 0.88113 | -0.19
Keitumetse 21-05 (1 hr) | 171 | 3.42 | 3.34 | 0.08 | 3.38 | 0.056569 | 1.6736 | -0.02
Keitumetse 21-05 (1 hr) | 172 | 3.24 | 3.29 | 0.05 | 3.265 | 0.035355 | 1.0829 | -0.135
Keitumetse 21-05 (1 hr) | 173 | 3.23 | 3.19 | 0.04 | 3.21 | 0.028284 | 0.88113 | -0.19
Keitumetse 21-05 (1 hr) | 174 | 3.17 | 3.22 | 0.05 | 3.195 | 0.035355 | 1.1066 | -0.205
Keitumetse 21-05 (1 hr) | 175 | 3.2 | 3.25 | 0.05 | 3.225 | 0.035355 | 1.0963 | -0.175
Keitumetse 21-05 (1 hr) | 176 | 3 | 2.99 | 0.01 | 2.995 | 0.0070711 | 0.2361 | -0.405
Keitumetse 21-05 (1 hr) | 177 | 3.12 | 3.09 | 0.03 | 3.105 | 0.021213 | 0.68319 | -0.295
Keitumetse 20-05 (1h30) | 158 | 3.59 | 3.59 | 0 | 3.59 | 0 | 0 | 0.19
Keitumetse 20-05 (1h30) | 159 | 3.46 | 3.39 | 0.07 | 3.425 | 0.049497 | 1.4452 | 0.025
Keitumetse 20-05 (1h30) | 160 | 3.47 | 3.48 | 0.01 | 3.475 | 0.0070711 | 0.20348 | 0.075
Keitumetse 20-05 (1h30) | 161 | 3.45 | 3.49 | 0.04 | 3.47 | 0.028284 | 0.81511 | 0.07
Keitumetse 20-05 (1h30) | 162 | 3.58 | 3.6 | 0.02 | 3.59 | 0.014142 | 0.39393 | 0.19
Keitumetse 20-05 (1h30) | 163 | 3.48 | 3.45 | 0.03 | 3.465 | 0.021213 | 0.61221 | 0.065
Keitumetse 20-05 (1h30) | 164 | 3.53 | 3.63 | 0.1 | 3.58 | 0.070711 | 1.9752 | 0.18
Keitumetse 20-05 (1h30) | 165 | 3.46 | 3.44 | 0.02 | 3.45 | 0.014142 | 0.40992 | 0.05
Keitumetse 20-05 (1h30) | 166 | 3.39 | 3.37 | 0.02 | 3.38 | 0.014142 | 0.41841 | -0.02
Keitumetse 20-05 (1h30) | 167 | 3.36 | 3.4 | 0.04 | 3.38 | 0.028284 | 0.83681 | -0.02
Lilo 13-06 (1 hr) | 1168 | 3.11 | 3.02 | 0.09 | 3.065 | 0.06364 | 2.0763 | -0.335
Lilo 13-06 (1 hr) | 1169 | 3.06 | 3.06 | 0 | 3.06 | 0 | 0 | -0.34
Lilo 13-06 (1 hr) | 1170 | 3.06 | 3.15 | 0.09 | 3.105 | 0.06364 | 2.0496 | -0.295
Lilo 13-06 (1 hr) | 1171 | 3.14 | 3.11 | 0.03 | 3.125 | 0.021213 | 0.67882 | -0.275
Lilo 13-06 (1 hr) | 1172 | 3.13 | 3.11 | 0.02 | 3.12 | 0.014142 | 0.45327 | -0.28
Lilo 13-06 (1 hr) | 1173 | 3.09 | 3.07 | 0.02 | 3.08 | 0.014142 | 0.45916 | -0.32
Lilo 13-06 (1 hr) | 1174 | 3.02 | 2.99 | 0.03 | 3.005 | 0.021213 | 0.70593 | -0.395
Lilo 13-06 (1 hr) | 1175 | 2.98 | 2.97 | 0.01 | 2.975 | 0.0070711 | 0.23768 | -0.425
Lilo 13-06 (1 hr) | 1176 | 3.05 | 3.05 | 0 | 3.05 | 0 | 0 | -0.35
Lilo 13-06 (1 hr) | 1177 | 3.12 | 3.05 | 0.07 | 3.085 | 0.049497 | 1.6045 | -0.315
Lilo 17-06 (1h30) | 1218 | 3.17 | 3.14 | 0.03 | 3.155 | 0.021213 | 0.67237 | -0.245
Lilo 17-06 (1h30) | 1219 | 2.89 | 2.98 | 0.09 | 2.935 | 0.06364 | 2.1683 | -0.465
Lilo 17-06 (1h30) | 1220 | 3.12 | 3.12 | 0 | 3.12 | 0 | 0 | -0.28
Lilo 17-06 (1h30) | 1221 | 3.07 | 3.03 | 0.04 | 3.05 | 0.028284 | 0.92735 | -0.35
Lilo 17-06 (1h30) | 1222 | 3.69 | 3.64 | 0.05 | 3.665 | 0.035355 | 0.96468 | 0.265
Lilo 17-06 (1h30) | 1223 | 3.03 | 2.97 | 0.06 | 3 | 0.042426 | 1.4142 | -0.4
Lilo 17-06 (1h30) | 1224 | 3 | 3.03 | 0.03 | 3.015 | 0.021213 | 0.70359 | -0.385
Lilo 17-06 (1h30) | 1225 | 2.89 | 2.96 | 0.07 | 2.925 | 0.049497 | 1.6922 | -0.475
Lilo 17-06 (1h30) | 1226 | 2.93 | 3.02 | 0.09 | 2.975 | 0.06364 | 2.1391 | -0.425
Lilo 17-06 (1h30) | 1227 | 2.9 | 2.97 | 0.07 | 2.935 | 0.049497 | 1.6865 | -0.465

## SUMMARY PER RUN
Run | n samples | mean of lab means | sd of lab means | mean within SD | mean bias | mean pctRSD
Keitumetse 20-05 (1h30) | 10 | 3.4805 | 0.080604 | 0.024749 | 0.0805 | 0.71102
Keitumetse 21-05 (1 hr) | 10 | 3.234 | 0.12814 | 0.028284 | -0.166 | 0.86919
Lilo 13-06 (1 hr) | 10 | 3.067 | 0.048028 | 0.025456 | -0.333 | 0.82652
Lilo 17-06 (1h30) | 10 | 3.0775 | 0.22061 | 0.037477 | -0.3225 | 1.2368

## Combined_metrics
overall mean | s_R | s_r (pooled repeatability sd) | r_limit | R_limit
3.2148 | 0.21374 | 0.035178 | 0.098499 | 0.59846