# Proficiency Testing (PT) Participation Plan (2025 – 2029)

Laboratory Name: Azuresky Pty Ltd

Document No: QMS-AS001 PT

Revision: 01

Effective Date: 13/04/2026

Next Review: 04/2030

Prepared by: Laboratory Manager		Signature: __________________

Approved by: Operations Manager		Signature: __________________

## 1. Objective

To define the laboratory’s 5-year strategy for participation in internal and external proficiency testing schemes. The plan ensures continuous verification of analytical competence for coal-related parameters under controlled, low-throughput conditions where PT complements other quality-assurance (QA) measures.

## 2. Scope

Parameter; Method/Standard. Quick Ash; In-house method (based on ISO 1171). ISO Ash; ISO 1171. ISO Volatile Matter; ISO 562. ISO Moisture (Inherent Moisture); SANS 5925. Calorific Value; ISO 1928. Total Moisture; ISO 589. Preparation; ISO 13909-4.

## 3. Laboratory Context

Operational focus: Quality assurance and validation of homogenised reference coal materials. Testing frequency is occasional, primarily for CRM stability checks and verification of homogenised lots. Duplicate analyses are performed for each batch, accompanied by CRMs to monitor drifts.

Internal QC practices include weekly verification of crushers, furnaces, and ovens (every Monday). Furnace and oven calibration is externally done by Micron, while the calorimeter is calibrated in-house. External technicians service crushers and pulverisers.

## 4. Quality-Assurance Measures Supporting PT

| QA / IQC Measure; Frequency; Purpose. CRM analysis with each duplicate batch; Each testing event; Detect and quantify drift or bias. Duplicate sample analysis; Each testing event; Check repeatability. Crusher verification; Weekly (Monday); Validate sample-preparation uniformity. Furnace & oven verification; Weekly (Monday); Confirm temperature accuracy. CRM stability check; Monthly; Detect degradation of stored reference materials. External calibration (Micron); Annual & every two years for mass pieces; Ensure metrological traceability. External servicing of mechanical units; As scheduled or when required; Maintain performance stability. |  |
| --- | --- |

## 5. Risk-Based Determination of PT Frequency

Risk Factor; Risk Level; Mitigation / QA Action. Limited sample throughput; Low; Monthly CRM verification plus quarterly PT participation. Staff turnover; Low; Cross-training and annual competency review. Instrument drift; Moderate; Duplicate,CRM monitoring, weekly verifications. Method complexity; Moderate; Periodic bias trending and external PT. External PT availability; Readily available; Enrolled quarterly.

## 6. Areas of Technical Competence

Technical Area; Measurement Process; Parameters Included; Key Equipment. Proximate Analysis; Gravimetric/ Thermogravimetric; Quick Ash, ISO Ash, ISO Volatile Matter, ISO Moisture; Muffle furnace, drying oven. Calorific-Value Determination; Combustion Calorimetry; Calorific Value; Bomb calorimeter. Total Moisture; Thermogravimetric; Total Moisture; Drying oven, Top Pan Balance. Sample Preparation; Crushing, Splitting & Milling; Preparation of Coal; Crushers, Splitters, Pulverizers.

## 7. PT Participation Strategy and Frequency

Parameter; External PT (Coal Concept); Internal PT (CRM Checks); Total Frequency; Rationale. Quick Ash; Quarterly; Monthly CRM check + duplicates; 4 external + 12 internal; Moderate risk; temperature-dependent. ISO Ash; Quarterly; Monthly CRM check; 4 external + 12 internal; Linked to calorific correlation. ISO Volatile Matter; Quarterly; Monthly CRM check; 4 external + 12 internal; Sensitive to furnace behaviour. ISO Moisture; Quarterly; Monthly CRM check; 4 external + 12 internal; Influenced by humidity. Calorific Value; Quarterly; Monthly CRM check; 4 external + 12 internal; Key coal quality indicator. Total Moisture; Bi-monthly; -; 6 per year; Moderate risk, supplier-provided frequency.. Sample Preparation; Bi-monthly; -; 6 per year; Moderate risk, supplier-provided frequency..

## 8. Graphical Summary – PT Cycle

The diagram below illustrates the PT cycle and review process:

Weekly Equipment Verification → Monthly CRM/Duplicate Check → Quarterly External PT (Coal Concept) → Bias & Trend Evaluation → Corrective Actions if Needed → Annual Management Review → PT Plan Update

## 9. Conclusion

This PT Plan demonstrates a sound, risk-based approach suited to a testing, proficiency testing provider & certified reference material producing laboratory with limited sample throughput. Through structured internal CRM verification, weekly equipment checks, quarterly & bi-monthly external PT participation, the laboratory ensures ongoing validation of analytical competence and readiness for SANAS accreditation compliance.

| Date | Revision no | Content Changed |
| --- | --- | --- |
| 13/04/2026 | 01 | No changes, only the following additions made: - PT plan scope increased to include PT participation for Total moisture & sample preparation - Areas of technical competence & PT participation section updated to include total moisture and sample preparation. - Conclusion updated to include bi-monthly pt participation of sample preparation & total moisture |