Production Plan

During the production planning of Proficiency Testing Scheme, homogeneity test is conducted (10 random samples are extracted from the sample/material product, tested for Quick Ash%). The samples are retained for determination of stability test for a period of 4 (four) weeks. The material retained after the PT production is packaged for Certified reference Material production. The goal of the production plan is to design the workflow that minimises wastage, attains PT and RM samples that have achieved sufficient homogeneity and stability requirements.

# General Information

One of the Certified Reference Material to be produced is defined in table below:

| Field | Entry |
| --- | --- |
| CRM Name | GA CRM AS002 |
| Unique batch Number | CRM AS002 2025-05 |
| Intended Use | Calibration, Control Sample, and Training of staff |
| Matrix type | Pulverised Coal |
| Properties to be assigned | ISO Ash, ISO Volatile Matter, Calorific Value, and Total Sulphur |
| Production Start date | 2025-04-29 |
| Production End Date | - |
| Responsible Person | Sifiso Nyambi |
| Approved by | Nonhlanhla Msibi |

# Materials and Equipment

A list of materials and equipment to be used in the production of the GA CRM AS002 are tabled and listed as follows:

NB: The supplier of the materials will vary due to the required CRM quality to be produced. However, it is required that a record of list of suppliers of coal associated which a specific CRM be maintained at the laboratory. This is to ensure that the same CRM will be produced that is consistent with the initial quality produced.

NB: All other equipment used for the preparation and testing, the validation or equipment maintenance are specified in the respective coal testing methods ie. Procedure on sample preparation of Hard coal has detailed how to use crushers and efficiency requirements needed to produce coal size material required.

| Item | Description/Specs | Supplier | UoM | Serial Number | Qualification Service |
| --- | --- | --- | --- | --- | --- |
| Raw Material | Bituminous coal sample | Elandsfontein Mine | - | - | Approved |
| Crushers | 25,10,3 (mm) | MegaSize | - | 202203331 SC, AS8809, A124119; respectively | Cleaned and verified |
| Pulveriser | 212 µm | MegaSize | - | 22L90713190645 | Cleaned and verified |
| Platform Scale | 0.01 Kg | Darvan Brenco | ±0.03 kg | 120330150 | Calibrated |
| Top Pan Balance | 0.1 g accuracy | AE Adam | ±0.3 g | AEA90242156 | Calibrated |
| IM Moisture Oven 50L | Maintain a temperature of 107oC | Darvan Brenco | ±1.0 oC | DOV53A/210325/40 | Calibrated |
| Mass pieces | 1 to 100 (g) | Micron | ±0.00014 | OB WS1.1R | Calibrated |
| Ash Furnace | Maintain 815±10oC for at least 1 hour | Micron (ATS) | ±4 oC | ATS/210325/7A | Calibrated |
| ISO Volatile Matter Furnace | 900oC±5oC | Micron (ATS) | ±3 oC | ATS/210325/59V | Calibrated |
| Packaging Material | ±150 g | Foamy | - | - | Cleaned |
| Homogeneity Drum | - | - | - | - | Cleaned |

# Processing Plan

| Step | Description | Equipment used | Environmental Conditions | Acceptance criteria |
| --- | --- | --- | --- | --- |
| 1 | Air drying | - | Atmospheric temperature | Sample should be visibly dry |
| 2 | Stage Crushing | 25,10,3 mm Crushers | Room temperature | Crusher Efficiency ≥90% |
| 3 | Pulverise | Serial no: 22L90713190645 | Room temperature | Pulveriser efficiency ≥95% |
| 4 | Homogeneity Drum | - | - | ¬ 4 hours |
| 5 | Homogeneity Validation (analysis) | ISO Ash Furnace | 19-26 oC | ANOVA, SSbetween<Check Value |
| 6 | Packaging | ±150g | Room temperature | - |
| 7 | Storage | Storage room | ≤ 20oC | Temperature <23oC |

# Homogeneity and Stability Strategy

| Assessment Type | Method | No of units | Parameter | Acceptance Criteria |
| --- | --- | --- | --- | --- |
| Homogeneity | HOM-01 | 10 | Quick Ash | ANOVA, SSbetween<Check Value |
| Short-term stability | STA-02 | 3 time points | Quick Ash | No trend or drift outside limits. |
| Long-term stability | STA-03 | 6 months, 12 months | Calorific Value and Q-ash | Deviation should not be >0.3MJ/kg and 3% of assigned value. |

NB: A monthly proficiency testing report has recorded the respective homogeneity and statibility test of respective CRM. A long term stability test record is maintained.

# Measurement Procedures

| Parameter | Method Reference | Equipment | Analyst(s) | Estimated UoM |
| --- | --- | --- | --- | --- |
| Q-Ash | ISO 1171 | Furnace | Keitumetse M |  |
| Volatile Matter | ISO 562 | Furnace | Lilo M |  |
| Calorific Value | ISO 1928 | Bomb Calorimeter | Nosipho H |  |
| Inherent Moisture | SANS 5925 | Drying Oven | Keitumetse M |  |

NB: The following authorised procedures are to be used:

- AS002IW Procedure on the conditioning of samples,

- AS003IW Procedure for the determination of Ash content

- AS006IW Procedure for the determination of Volatile Matter

- AS004IW Procedure on the determination of Inherent Moisture

- AS007IW Procedure on the Calorific Value determination.

# Risk Assessment

| Risk Description | Impact | Mitigation strategy |
| --- | --- | --- |
| Power outage during analysis | Low | Backup generator |
| Balance drift | Medium | Verification with calibrated mass pieces at the beginning of every shift. |
| Sample Swap | High | Prepare one sample at a time, only after packaging has been completed then the next sample can be prepared. |

# Records and Approval

| Record Type | File Name/ Location | Responsible Person | Retention Period |
| --- | --- | --- | --- |
| Batch Processing | Form | Laboratory Supervisor | 5 years |
| Analytical Results Sheet | LIMS | Laboratory Supervisor | 5 years |
| Homogeneity Test report and Stability Test report | [can be retrieved from within the final PT report] or website | Laboratory Manager/Supervisor | 10 years |
| Calibration Certificates | Technical File | Lab Manager | 5 Years |