Kindly note  that we have two different thermometer and probes. So, the following should be considered when doing the verifications:

- The thermometer and probe come as a set: For example, the thermometer is 1b and the probe is 1a, together they form 1a+b. These should be used as a unit. At no point should you use 1a with 2b.

- Take note of the correction factors:

| OBS T1.1a+b | 40oC | 105oC | 500oC | 815oC | 900oC |
| --- | --- | --- | --- | --- | --- |
| Correction Factor (oC) | 0 | -1.8 | -3 | -5 | -3 |

| OBS T2.1a+b | 40oC | 105oC | 500oC | 815oC | 900oC |
| --- | --- | --- | --- | --- | --- |
| Correction Factor (oC) | 0.7 | -0.6 | -1 | 0 | 3 |

NB: Although there is a negative sign next to the correction factors, the principle is to add (i.e. we always add the correction factor).

Certified Reference Material used for drift checks is CRM 002 and the values are as following:

| Measurand | Assigned Value | Uncertainty (±) | Unit |  | Method |
| --- | --- | --- | --- | --- | --- |
| ISO Ash | 23.33 | ± 0.21 | % |  | ISO 1171 |
| ISO Volatile Matter | 26.13 | ± 0.60 | % |  | ISO 562 |
| Calorific Value | 24.92 | ± 0.15 | MJ/kg |  | ISO 1928 |
| Total Sulphur | 1.46 | ± 0.03 | % |  | ASTM D4239 |
| Quick Ash | 23.29 | ± 0.271 | % |  | Internal Method |

| Measurand | Value | Tolerance | Unit | Minimum | Maximum | Method |
| --- | --- | --- | --- | --- | --- | --- |
| ISO Ash | 22.54 | 0.68 | % | 21.86 | 23.21 | ISO 1171 |
| ISO Volatile Matter | 25.24 | 1.01 | % | 24.23 | 26.25 | ISO 562 |
| Calorific Value | 24.07 | 0.30 | MJ/kg | 23.77 | 24.37 | ISO 1928 |
| Total Sulphur | 1.43 | 0.15 | % | 1.28 | 1.58 | ASTM D4239 |
| Quick Ash | 22.50 | 0.67 | % | 21.83 | 23.17 | Internal Method |