Characterisation

Objectives

To assign a property value to a reference material using scientifically valid, traceable, and statistically sound techniques.

# Scope

This procedure describes how to determine the certified values of reference materials (RMs/CRMs) through characterisation in compliance with ISO 17034:2016 Clause 7.12.

# Reference Standards

- ISO 17034:2016 — General requirements for the competence of RMPs

- ISO/IEC 17025 — General requirements for the competence of testing and calibration laboratories

- ISO 33405:2024— Reference materials – Guidance for characterisation and assessment of homogeneity and stability

- Specific test method standards (e.g. ISO 1171, ISO 562, ISO 589, etc.)

# Responsibilities

- Analyst: Conducts characterisation measurements

- Laboratory Manager: Process data, calculate assigned values, reviews and approves characterisation reports

- Operations Manager: Reviews and approves all procedures, technical forms.

# Characterisation Strategy

The lab will be using the following approaches:

- One laboratory using validated methods

- Multiple laboratories using the same or different validated methods

- Combination of methods or historical data, where justified.

# Procedure step

#### 5.1. Define the Measurand

Clearly state the property (e.g., Ash content in % on dry basis) and units.

#### 5.2. Select the Measurement Method

- Use methods validated per ISO/IEC 17025 or ISO 5725.

- The method must be appropriate for the measurand and matrix.

- Document all method parameters (temperature, mass, sample prep, etc.).

#### 5.3. Ensure Metrological Traceability

- Use instruments calibrated with standards traceable to the SI units.

- Use CRM reference materials if applicable (e.g., for calorific value).

#### 5.4. Perform the Measurements

- Use at least 2 replicate measurements per unit to ensure repeatability.

- If multiple units are used, select them randomly.

- If multiple labs are involved, standardise procedures and sample prep.

#### 5.5. Control Environmental Conditions

Ensure environmental conditions (e.g., temperature, humidity) are suitable for accurate and stable measurements.

#### 5.6. Evaluate and Document Raw Data

- Record all instrument settings, analyst IDs, and time/date.

- Secure all records in a traceable, audit-proof system.

#### 5.7. Statistical Treatment

- Use means or robust statistical estimators (e.g., robust mean, median).

- Identify and eliminate outliers using ISO 5725, Dixon’s Q, or Grubbs' test or any other sound statistical means.

- Calculate the standard deviation of the mean or robust uncertainty.

#### 5.8. Assess Method Uncertainty

- Include method repeatability, bias, calibration, and sample preparation effects.

- Uncertainty should be documented in an uncertainty budget.

#### 5.9. Assign Property Values

- The assigned value is the best estimate of the true value based on valid statistical analysis.

- Include expanded uncertainty (U = k × u), coverage factor (usually k=2 for 95% confidence).

#### 5.10. Peer Review / Technical Review

- Independent review by technical staff not involved in measurement.

- Check method validation, traceability, uncertainty, and data integrity.

# Documentation and Reporting

The following should be included in the characterisation report:

- RM Name, ID, Batch

- Measurand and method used

- Measurement results

- Assigned value and uncertainty

- Traceability chain

- Analysts, labs involved

- Statistical treatment

The records should be archived for a period of 5 years.

# Controls and Records

| Record | Description |
| --- | --- |
| Raw data sheets | Original measurement results |
| Calibration certificates | For all measurement instruments used |
| Uncertainty budget | Used to estimate uncertainty |
| Statistical reports | All analysis used to determine certified values |
| Characterisation reports | Final signed report for each batch |