Metrological Traceability

Objective

To ensure that all certified values assigned to reference materials are metrologically traceable to the international system of units (SI units).

# Requirements

- Certified values shall be traceable to the SI unit or other recognised accepted standard.

- Should SI traceability not be possible, the laboratory shall use other technically justifiable references such as consensus values, comparison with CRMs.

- The laboratory shall document and justify traceability route and maintain evidence to support it.

# Establishing traceability route

| Parameter | Unit | Required traceability to | Standard/reference |
| --- | --- | --- | --- |
| Calorific Value | MJ/kg | SI (Joules) via CRM (benzoic acid) | Benzoic acid tablet from supplier |
| Ash Content | % | SI via grams (g) | Calibrated balances traceable to SI |
| Moisture | % | SI- Mass difference | Traceable mass calibration |
| Volatile | % | SI- Mass and temperature | Calibrated oven and balance |
| Total Sulphur | % | SI via grams (g) | - |

# Required elements of traceability system for the laboratory

- Use of traceable standards and instruments

- Use CRM: e.g. benzoic acid tablets (bomb calorimetry)

- Calibrate instruments using services from Mass & Temperature metrology ISO/IEC 17025 accredited laboratory’s: Balances, thermometers, ovens, bomb calorimetry.

- Validation

- Ensure that all methods (ISO 1928, ISO 1171 etc) are validated in-house and that all standards are traceable.

- Establishing the traceability chain

e.g. Benzoic acid CRM (with CoA stating SI traceability) → Calorimeter → Coal sample → Assigned value

# Traceability statements

For e.g. :

- “The assigned value for Gross Calorific Value is traceable to the International System of Units (SI) through calibration with certified benzoic acid, supplied by an ISO 17034-accredited producer. All equipment used is calibrated by ISO/IEC 17025-accredited laboratories”.

- “The certified value for Ash Content is traceable to the SI unit of mass (grams) through gravimetric determination using balances calibrated by ISO/IEC 17025-accredited laboratories. The method applied is based on ISO 1171, and the results are supported by the use of traceable mass standards and validated furnace temperature control”.

- "The certified value for Volatile Matter is traceable to the SI units of mass (grams) and temperature (Kelvin), using calibrated analytical balances and high-temperature ovens verified by ISO/IEC 17025-accredited laboratories. The determination method complies with ISO 562. Traceability is maintained via certified reference materials and controlled thermal decomposition conditions."

- “The certified value for Inherent Moisture is traceable to the SI unit of mass (grams) through mass loss determination using a balance calibrated by an ISO/IEC 17025-accredited laboratory. The method used is in accordance with ISO 589, and traceability is supported by certified mass standards and validated drying procedures with controlled temperature and time settings."

- "The certified value for Total Sulphur is traceable to the SI unit of mass fraction (mg/kg) through calibration with certified reference materials supplied by ISO 17034-accredited producers. Measurement instrumentation is calibrated and verified by ISO/IEC 17025-accredited laboratories, and determinations are performed using validated combustion or IR detection methods in accordance with ISO 19579 (or ASTM D4239)."