## 1. Purpose

To define the systematic approach used by the laboratory to monitor the validity of test results and ensure continuing accuracy, reliability, and traceability in compliance with ISO/IEC 17025:2017, Clause 7.7.

## 2. Scope

This procedure applies to all analytical preparation and determinations performed by the laboratory, including Preparation & Proximate Analysis (Moisture, Ash, Volatile Matter, and Fixed Carbon) and Calorific Value (CV) testing, as well as to all instruments and reference materials associated with these determinations.

## 3. Responsibilities

| Role | Responsibilities |
| --- | --- |
| Laboratory Manager / Supervisor | Approves the monitoring programme, reviews trends, and authorizes corrective actions. Ensure adherence to monitoring frequency, maintain QC logs, verify control chart performance, and review duplicates and CRMs. |
| Laboratory Technician | Perform verification checks, run duplicates and CRMs, record results accurately, and initiate investigation for out-of-control data. |
| Operations Manager | Reviews all records for completeness and trends, ensures calibration certificates are current, and reports performance in management review. |

## 4. Definitions

- CRM: Certified Reference Material — a material with a certified analyte value and associated uncertainty.

- Control Chart: Graphical tool used to monitor performance trends and detect bias or drift.

- Repeatability Limit: Maximum allowable difference between duplicate results under the same conditions.

- Verification: Confirmation that an instrument performs within defined limits before use.

## 5. Monitoring Methods

## 5.1 Instrument Calibration and Verification

- Instruments that require calibration — Ash Furnace, Volatile Furnace, Moisture Oven, Drying Oven, Analytical Balances, and Mass Pieces — are sent to SANAS-accredited calibration facilities to ensure traceability to national standards.

- Calibration certificates are reviewed upon receipt and retained in the instrument calibration file.

- Verification prior to use:

  - Analysts perform daily or batch verification checks (e.g., temperature control and recording, balance verification, equipment efficiency) before starting work.

  - Any deviations beyond control limits are reported to the Laboratory Manager/supervisor before testing proceeds.

- CV instrument (Calorimeter) is calibrated in-house using a Certified Benzoic Acid pill of known calorific value. Calibration results are documented and plotted on a control chart to assess stability.

## 5.2 Analytical Quality Control

- Duplicate Analyses:

  - Each sample run includes at least one duplicate test sample to evaluate precision.

  - Duplicates must meet the following repeatability limits:

| Parameter | Repeatability Limit |
| --- | --- |
| Calorific Value (CV) | ≤ 0.12 MJ/kg |
| Ash (≤ 10 %) | ≤ 0.2 % abs |
| Ash (> 10 %) | ≤ 2 % of mean |
| Volatile (≤ 10 %) | ≤ 0.3 % abs |
| Volatile (> 10 %) | ≤ 3 % of mean |
| Inherent Moisture (≤ 5 %) | ≤ 0.1 % abs |
| Inherent Moisture (> 5 %) | ≤ 0.15 % abs |

  - Results exceeding limits must be re-run, if it still is out of repeatability limit, run a CRM. If problem persists, do a quick verification/maintenance on instrument. Conduct an investigation if there is any drift found on verification/maintenance.

- Certified Reference Material (CRM):

  - A CRM is analysed after every 10 routine runs for each parameter.

  - The measured value is compared with the certified value; deviations beyond control limits (reproducibility) are investigated.

- Control Charts:

  - Control charts are maintained for each parameter using CRM and duplicate data.

  - Each new result is plotted and reviewed to detect bias, drift, or loss of precision.

  - Out-of-control points (beyond Reproducibility limits) or systematic trends require immediate suspension of affected runs until resolution.

## 5.3 Method Validation and Ongoing Verification

- All methods are validated before routine use to confirm accuracy, precision, linearity, and detection limits.

- Periodic re-validation is performed after any significant change (New equipment).

- Analysts verify method performance using QC data trends and inter-analyst comparisons.

## 5.4 Proficiency Testing and Interlaboratory Comparison

- The laboratory participates in proficiency testing (PT) for each accredited parameter.

- PT performance is evaluated using z-scores and included in management review.

- If no PT is available, inter-laboratory comparisons are conducted

5.5 Trending and Review

- All QC data (duplicates, CRMs, calibrations, PT results) are trended per control chart.

- The Laboratory Manager and Laboratory Supervisor jointly review control charts.

- Any observed bias, drift, or degradation of performance triggers a formal Investigation.

- Persistent trends may result in method re-validation, recalibration, or staff retraining.

## 6. Corrective Action Procedure

- Stop analysis of samples for affected parameter or batch.

- Re-verify instrument verification, calibrate where necessary and perform impact analysis on affected batches.

- Document root cause, corrective and preventive action in the Nonconformance Register.

- Confirm restored control with CRM or duplicate results before resuming routine testing.

## 7. Records and Documentation

The following records are retained electronically or in bound QC logbooks:

| Record | Retention Period | Responsibility |
| --- | --- | --- |
| Calibration Certificates | 5 years | Operations Manager |
| Daily Verification Logs | 3 years | Analyst |
| CRM Results & Control Charts | 5 years | Supervisor |
| Duplicate Test Records | 3 years | Analyst |
| PT/Inter-lab Results | 5 years | Operations Manager |
| IRQs and Corrective Actions | 5 years | Quality Manager |

## 8. Review

This procedure shall be reviewed annually during management review meetings.

## 9. References

- ISO/IEC 17025:2017 — Clauses 7.6, 7.7, 7.8

- ILAC G8:09/2019 — Guidelines on decision rules and reporting

- ASTM D3173, D3174, D3175, D5865 — Coal analysis standards

- Laboratory QMS Manual, Section 7 — Quality Control and Validation

10. Revision

| Revision no | Content changed | Reason for revision | Date changed |
| --- | --- | --- | --- |