## Verification Procedure
7. Sieve Verification (ISO 3310-1 / ISO 3310-2)
To maintain the accuracy of the PSD method, all sieves must be verified every six months.
7.1. Why We Verify
Coal is highly abrasive. Over time, the wires in the mesh can shift or wear down (becoming thinner), which effectively increases the size of the holes. If a 10mm sieve wears down to 10.5mm, your PSD results will be biased toward the smaller fractions.
7.2. Measurement Protocol
1. Cleaning: Ensure the sieve is manually brushed to remove all "blinded" particles.
2. Sampling the Mesh: You cannot measure every hole. You must measure 10 apertures chosen along a diagonal line (45°) across the mesh(for the smaller aperture screens).
3. Two-Way Measurement: For each of the holes, use the internal jaws of a calibrated Digital Vernier Caliper to measure:
Dimension X: The horizontal width of the opening.
Dimension Y: The vertical height of the opening.
Note: Ensure the jaws are flush against the wires but not forcing them apart.
7.3. Calculating the Result
1. Mean Aperture (w): Calculate the average of all 20 measurements (10X and 10Y).
2. Tolerance Check: Compare the Mean Aperture against the ISO 3310-1 Tolerance Table.
Example: For a 6.30 mm sieve, the typical tolerance is roughly 0.2 mm.
If your average is 6.55 mm, the sieve has stretched beyond the limit.
7.4. Final Documentation
If the sieve passes, update the Sieve Calibration Log and apply a new green "Verified" sticker.
If the sieve fails, it must be etched with "REJECTED" or the mesh must be physically punctured to prevent accidental use in the lab.
7. Verification of Large Apertures (> 12.5 mm)
Visual Check: Ensure the plate is flat. Large screens often "bow" or warp under the weight of 65 kg of coal. A warped screen changes the effective aperture size and must be replaced.
Measurement: Use a calibrated Vernier Caliper to measure the internal distance of 10 square apertures across the plate.
Bridge Inspection: Inspect the metal strips between the apertures for thinning. If the metal bridge has worn down by more than 25% of its original thickness, the sieve must be decommissioned, as it will likely flex and bypass larger particles during the "8 revolutions" screening process.
Sieve Type | Size Range | Min. Apertures | Total Measurements | Pattern
Woven Wire | < 12.5 mm | 10 | 20 (X & Y) | Diagonal Line
Perforated Plate | > 12.5 mm | 10 | 20 (X & Y) | "X" or Cross Pattern

## Tolerances
1. Fine & Medium Sieves (ISO 3310-1: Woven Wire)
Table of tolerance of sieve apetures( smaller apeture screen)
Nominal Aperture (w) | Tolerance on Average (±Y) | Max Individual Aperture (X) | Standard Deviation (σ)
11.20 mm | 0.360 mm | 12.10 mm | 0.690 mm
8.00 mm | 0.260 mm | 8.71 mm | 0.490 mm
5.60 mm | 0.190 mm | 6.16 mm | 0.350 mm
2.80 mm | 0.110 mm | 3.15 mm | 0.190 mm
1.00 mm | 0.040 mm | 1.17 mm | 0.080 mm
212µm (0.212 mm) | 0.011 mm | 0.263 mm | 0.024 mm
Average Tolerance (Y): If the average of your 20 measurements is outside this range, the sieve is out of calibration.
Max Aperture (X): If any single hole measures larger than this, the sieve is REJECTED immediately.
2. Coarse Sieves (ISO 3310-2: Perforated Plate)
These are usually square-hole or round-hole punched plates.
Nominal Aperture (w) | Tolerance on Average (±Y) | Max Individual Aperture (X) | Min. Plate Thickness
50.0 mm | 1.50 mm | 52.40 mm | 3.0 mm
31.5 mm | 1.00 mm | 33.20 mm | 2.5 mm
22.4 mm | 0.70 mm | 23.70 mm | 2.0 mm
16.0 mm | 0.50 mm | 17.10 mm | 1.5 mm
Plate Thickness: For larger screens, also measure the thickness of the metal. If the plate has worn thin due to the abrasive nature of coal, it will flex and allow "oversize" particles to pass during the 360° revolution.

## Summarized Values
Doc No: PSD014
Compiled By: TMJ Maponya
Approved By: SB Nyambi
Effective Date: 15 April 2026
3. Overall Sieve Verification Log
Sieve ID | Nominal Size | Date | Average | Max Found | Status (Pass/Fail) | Verified By
500401 | 1mm | 2026-04-17 | 1.0035 | 1.03 | PASS | Tshepo Maponya
500623 | 3mm | 2026-04-17 | 3.0015 | 3.05 | PASS | Tshepo Maponya
500628 | 4,75mm | 2026-04-17 | 4.751 | 4.81 | PASS | Tshepo Maponya
500847 | 6mm | 2026-04-17 | 6.005 | 6.06 | PASS | Tshepo Maponya
600311 | 10mm | 2026-04-17 | 12.001 | 12.2 | PASS | Tshepo Maponya
600519 | 25mm | 2026-04-17 | 25.02 | 25.3 | PASS | Tshepo Maponya
610209 | 22.4mm | 2026-04-17 | 22.2235 | 22.5 | PASS | Tshepo Maponya
500697 | 50mm | 2026-04-17 | 50.035 | 50.4 | PASS | Tshepo Maponya
Vernier Caliper Used : | 0-150 Digital Caliper
Certificate Number: | DB-14965
UoM: | ±0,015mm
Serial Number: | MG-10

## Verification
Aperture #(Sieve 1 mm) | Dimension X (mm) | Dimension Y (mm) |  |  | Aperture #(Sieve 3 mm) | Dimension X (mm) | Dimension Y (mm) | Aperture #(Sieve 4,75 mm) | Dimension X (mm) | Dimension Y (mm) | Aperture #(Sieve 6 mm) | Dimension X (mm) | Dimension Y (mm) | Aperture #(Sieve 12 mm) | Dimension X (mm) | Dimension Y (mm) | Aperture #(Sieve 25 mm) | Dimension X (mm) | Dimension Y (mm) | Aperture #(Sieve 50 mm) | Dimension X (mm) | Dimension Y (mm) | Aperture #(Sieve 22,4 mm) | Dimension X (mm) | Dimension Y (mm)
1 | 1.01 | 0.99 |  |  | 1 | 3.02 | 2.98 | 1 | 4.77 | 4.76 | 1 | 6.05 | 5.95 | 1 | 12.1 | 11.9 | 1 | 25.1 | 24.9 | 1 | 50.2 | 49.8 | 1 | 22 | 22.1
2 | 1 | 1.02 |  |  | 2 | 2.99 | 3.01 | 2 | 4.75 | 4.73 | 2 | 6 | 6.02 | 2 | 12 | 12.05 | 2 | 25.05 | 25 | 2 | 50.1 | 50.05 | 2 | 22.1 | 22.15
3 | 0.98 | 1.01 |  |  | 3 | 3.05 | 2.95 | 3 | 4.8 | 4.7 | 3 | 5.98 | 6.01 | 3 | 11.95 | 12.02 | 3 | 24.95 | 25.15 | 3 | 49.9 | 50.15 | 3 | 22.3 | 22.2
4 | 1.03 | 1 |  |  | 4 | 3 | 3 | 4 | 4.74 | 4.74 | 4 | 6.03 | 5.99 | 4 | 12.15 | 11.85 | 4 | 25.2 | 24.8 | 4 | 50.3 | 49.7 | 4 | 22.4 | 22.4
5 | 0.99 | 1.02 |  |  | 5 | 3.01 | 2.99 | 5 | 4.76 | 4.75 | 5 | 6.01 | 6 | 5 | 12.05 | 11.95 | 5 | 25 | 25.05 | 5 | 50 | 50.1 | 5 | 22.15 | 22
6 | 1.01 | 0.98 |  |  | 6 | 2.97 | 3.03 | 6 | 4.72 | 4.78 | 6 | 5.97 | 6.04 | 6 | 11.9 | 12.1 | 6 | 24.9 | 25.1 | 6 | 49.85 | 50.25 | 6 | 22.05 | 22.5
7 | 1.02 | 1.01 |  |  | 7 | 3.02 | 3.01 | 7 | 4.79 | 4.71 | 7 | 6.02 | 6.02 | 7 | 12.12 | 11.88 | 7 | 25.15 | 24.95 | 7 | 50.15 | 49.95 | 7 | 22.17 | 22.4
8 | 1 | 1 |  |  | 8 | 2.98 | 3 | 8 | 4.75 | 4.75 | 8 | 6 | 6 | 8 | 12 | 12 | 8 | 25 | 25 | 8 | 50 | 50 | 8 | 22.2 | 22.1
9 | 0.97 | 1.03 |  |  | 9 | 3.04 | 2.96 | 9 | 4.81 | 4.69 | 9 | 6.06 | 5.94 | 9 | 12.2 | 11.8 | 9 | 25.3 | 24.7 | 9 | 50.4 | 49.6 | 9 | 22.4 | 22.4
10 | 1.01 | 0.99 |  |  | 10 | 3 | 3.02 | 10 | 4.75 | 4.77 | 10 | 6.01 | 6 | 10 | 12.05 | 11.95 | 10 | 25.1 | 25 | 10 | 50.2 | 50 | 10 | 22.3 | 22.15
COMBINED TOTAL | 1.0035 | Within limits |  |  | COMBINED TOTAL | 3.0015 | Within limits | COMBINED TOTAL | 4.751 | Within limits | COMBINED TOTAL | 6.005 | Within limits | Combined Total | 12.001 | Within Limits | Combined Total | 25.02 | Within Limits | TOTALS | 50.035 | Within Limits | Combined Total | 22.2235 | Within Limits
Std Deviation | 0.016631 |  |  |  | Std Deviation | 0.025603 |  | Std Deviation | 0.031439 |  | Std Deviation | 0.03 |  | Std Deviation | 0.10493 |  | Std Deviation | 0.13707 |  | Std Deviation | 0.20007 |  | Std Deviation | 0.15215
ISO Limit (X) | 1,17 mm |  |  |  | ISO Limit (X) | 3,37 mm |  | ISO Limit (X) | 5,26 mm |  | ISO Limit (X) | 6.58 mm |  | ISO Limit (X) | 12,90mm |  | ISO Limit (X) | 26,40mm |  | ISO Limit (X) | 52,40mm |  | ISO Limit (X) | 23,70mm
| Nominal Aperture (w) | Sieve Type | Avg. Tolerance (±Y) | Target Std Dev (σ) | Max Individual (X)
| 50.0 mm | Perforated | 1.50 mm | 1.80 mm | 52.40 mm
| 25.0 mm | Perforated | 0.80 mm | 1.00 mm | 26.40 mm
| 22,4 mm | Perforated | 0,70 mm | 0,90 mm | 23,70mm
| 12.0 mm | Woven/Perf | 0.40 mm | 0.70 mm | 12.90 mm
| 6.0 mm | Woven Wire | 0.20 mm | 0.38 mm | 6.58 mm
| 4.75 mm | Woven Wire | 0.16 mm | 0.30 mm | 5.26 mm
| 3.0 mm | Woven Wire | 0.11 mm | 0.19 mm | 3.37 mm
| 1.0 mm | Woven Wire | 0.04 mm | 0.08 mm | 1.17 mm