# Method Validation Report for Inherent Moisture (SANS 5925)

Method: SANS 5925:2007

Laboratory: Azuresky Pty Ltd

Reference value (CRM002): 3.40%

Analysts & dates: Keitumetse (21-05-2025, 20-05-2025); Lilo (13-06-2025, 17-06-2025)

Prepared on: 2025-07-22

## Introduction

This validation study was performed to confirm the suitability of the laboratory method for the determination of inherent moisture in coal samples, in accordance with SANS 5925. The validation covers repeatability, reproducibility, accuracy (trueness), and bias evaluation.

## Experimental Design

The method was validated through replicate analyses performed by two analysts (Keitumetse and Lilo) on different days. Each analyst performed two runs of 10 paired replicate measurements (R1 and R2). For each sample, the mean, within-sample standard deviation, repeatability (%RSD), and bias against the CRM reference value (3.40%) were calculated.

## Raw replicated data and per-sample calculations

Below are the raw replicate results (R1, R2), differences, sample means, within-sample SD, %RSD and bias vs CRM for each run and sample.

### Keitumetse 21-05 (1 hr)

| LIMS | R1 | R2 | Difference | Mean | Within SD | %RSD | Bias vs CRM |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 168 | 3.330 | 3.300 | 0.030 | 3.315000 | 0.021213 | 0.640% | -0.085000 |
| 169 | 3.430 | 3.450 | 0.020 | 3.440000 | 0.014142 | 0.411% | 0.040000 |
| 170 | 3.190 | 3.230 | 0.040 | 3.210000 | 0.028284 | 0.881% | -0.190000 |
| 171 | 3.420 | 3.340 | 0.080 | 3.380000 | 0.056569 | 1.674% | -0.020000 |
| 172 | 3.240 | 3.290 | 0.050 | 3.265000 | 0.035355 | 1.083% | -0.135000 |
| 173 | 3.230 | 3.190 | 0.040 | 3.210000 | 0.028284 | 0.881% | -0.190000 |
| 174 | 3.170 | 3.220 | 0.050 | 3.195000 | 0.035355 | 1.107% | -0.205000 |
| 175 | 3.200 | 3.250 | 0.050 | 3.225000 | 0.035355 | 1.096% | -0.175000 |
| 176 | 3.000 | 2.990 | 0.010 | 2.995000 | 0.007071 | 0.236% | -0.405000 |
| 177 | 3.120 | 3.090 | 0.030 | 3.105000 | 0.021213 | 0.683% | -0.295000 |

### Keitumetse 20-05 (1h30)

| LIMS | R1 | R2 | Difference | Mean | Within SD | %RSD | Bias vs CRM |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 158 | 3.590 | 3.590 | 0.000 | 3.590000 | 0.000000 | 0.000% | 0.190000 |
| 159 | 3.460 | 3.390 | 0.070 | 3.425000 | 0.049497 | 1.445% | 0.025000 |
| 160 | 3.470 | 3.480 | 0.010 | 3.475000 | 0.007071 | 0.203% | 0.075000 |
| 161 | 3.450 | 3.490 | 0.040 | 3.470000 | 0.028284 | 0.815% | 0.070000 |
| 162 | 3.580 | 3.600 | 0.020 | 3.590000 | 0.014142 | 0.394% | 0.190000 |
| 163 | 3.480 | 3.450 | 0.030 | 3.465000 | 0.021213 | 0.612% | 0.065000 |
| 164 | 3.530 | 3.630 | 0.100 | 3.580000 | 0.070711 | 1.975% | 0.180000 |
| 165 | 3.460 | 3.440 | 0.020 | 3.450000 | 0.014142 | 0.410% | 0.050000 |
| 166 | 3.390 | 3.370 | 0.020 | 3.380000 | 0.014142 | 0.418% | -0.020000 |
| 167 | 3.360 | 3.400 | 0.040 | 3.380000 | 0.028284 | 0.837% | -0.020000 |

### Lilo 13-06 (1 hr)

| LIMS | R1 | R2 | Difference | Mean | Within SD | %RSD | Bias Vs CRM |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1168 | 3.110 | 3.020 | 0.090 | 3.065000 | 0.063640 | 2.076% | -0.335000 |
| 1169 | 3.060 | 3.060 | 0.000 | 3.060000 | 0.000000 | 0.000% | -0.340000 |
| 1170 | 3.060 | 3.150 | 0.090 | 3.105000 | 0.063640 | 2.050% | -0.295000 |
| 1171 | 3.140 | 3.110 | 0.030 | 3.125000 | 0.021213 | 0.679% | -0.275000 |
| 1172 | 3.130 | 3.110 | 0.020 | 3.120000 | 0.014142 | 0.453% | -0.280000 |
| 1173 | 3.090 | 3.070 | 0.020 | 3.080000 | 0.014142 | 0.459% | -0.320000 |
| 1174 | 3.020 | 2.990 | 0.030 | 3.005000 | 0.021213 | 0.706% | -0.395000 |
| 1175 | 2.980 | 2.970 | 0.010 | 2.975000 | 0.007071 | 0.238% | -0.425000 |
| 1176 | 3.050 | 3.050 | 0.000 | 3.050000 | 0.000000 | 0.000% | -0.350000 |
| 1177 | 3.120 | 3.050 | 0.070 | 3.085000 | 0.049497 | 1.604% | -0.315000 |

### Lilo 17-06 (1h30)

| LIMS | R1 | R2 | Difference | Mean | Within SD | %RSD | Bias_vs_CRM |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1218 | 3.170 | 3.140 | 0.030 | 3.155000 | 0.021213 | 0.672% | -0.245000 |
| 1219 | 2.890 | 2.980 | 0.090 | 2.935000 | 0.063640 | 2.168% | -0.465000 |
| 1220 | 3.120 | 3.120 | 0.000 | 3.120000 | 0.000000 | 0.000% | -0.280000 |
| 1221 | 3.070 | 3.030 | 0.040 | 3.050000 | 0.028284 | 0.927% | -0.350000 |
| 1222 | 3.690 | 3.640 | 0.050 | 3.665000 | 0.035355 | 0.965% | 0.265000 |
| 1223 | 3.030 | 2.970 | 0.060 | 3.000000 | 0.042426 | 1.414% | -0.400000 |
| 1224 | 3.000 | 3.030 | 0.030 | 3.015000 | 0.021213 | 0.704% | -0.385000 |
| 1225 | 2.890 | 2.960 | 0.070 | 2.925000 | 0.049497 | 1.692% | -0.475000 |
| 1226 | 2.930 | 3.020 | 0.090 | 2.975000 | 0.063640 | 2.139% | -0.425000 |
| 1227 | 2.900 | 2.970 | 0.070 | 2.935000 | 0.049497 | 1.686% | -0.465000 |

## Summary statistics per run

| Run | n samples | Mean of lab means | SD of lab means | Mean within SD | Mean bias (vs CRM) |
| --- | --- | --- | --- | --- | --- |
| Keitumetse 20-05 (1h30) | 10 | 3.480500 | 0.080604 | 0.024749 | 0.080500 |
| Keitumetse 21-05 (1 hr) | 10 | 3.234000 | 0.128145 | 0.028284 | -0.166000 |
| Lilo 13-06 (1 hr) | 10 | 3.067000 | 0.048028 | 0.025456 | -0.333000 |
| Lilo 17-06 (1h30) | 10 | 3.077500 | 0.220608 | 0.037477 | -0.322500 |

Combined: n = 40, overall mean = 3.214750, reproducibility SD (s_R) = 0.213736

## Statistical calculations

Pooled repeatability standard deviation (s_r) calculated from the 40 within-run SDs: s_r = 0.035178

Repeatability limit r = 2.8 * s_r = 0.098499

Reproducibility limit R = 2.8 * s_R = 0.598462

## ANOVA (one-way) results

SSB = 1.116613, df_between = 3, MSB = 0.372204

SSW = 0.665035, df_within = 36, MSW = 0.018473

F-statistics = 20.148338

ANOVA p-value = 7.772889e-08

There is a statistically significant difference between runs (p < 0.05).

## Trueness / Bias t-tests (one-sample)

Run: Keitumetse 20-05 (1h30)

n = 10, mean = 3.480500, sd = 0.080604

Bias (mean - CRM) = 0.080500

t-statistic = 3.158212, two-sided p-value = 1.158348e-02

Interpretation: bias is significant at alpha = 0.05

Run: Keitumetse 21-05 (1 hr)

n = 10, mean = 3.234000, sd = 0.128145

Bias (mean - CRM) = -0.166000

t-statistic = -4.096442, two-sided p-value = 2.690982e-03

Interpretation: bias is significant at alpha = 0.05

Run: Lilo 13-06 (1 hr)

n = 10, mean = 3.067000, sd = 0.048028

Bias (mean - CRM) = -0.333000

t-statistics = -21.925617, two-sided p-value = 4.030381e-09

Interpretation: bias is significant at alpha = 0.05

Run: Lilo 17-06 (1h30)

n = 10, mean = 3.077500, sd = 0.220608

Bias (mean - CRM) = -0.322500

t-statistic = -4.622828, two-sided p-value = 1.249325e-03

Interpretation: bias is significant at alpha = 0.05

### Combined test (all lab means vs CRM)

Combined mean = 3.214750, n = 40, sd = 0.213736

t-statistic = -5.481629, two-sided p-value = 2.714753e-06

Interpretation: combined bias is significant at alpha = 0.05

## Measurement uncertainty (top-down)

Standard uncertainty u = s_R / sqrt(N) = 0.033795

Expanded uncertainty (k=2) U = 0.067589

## Plots

Boxplot of lab means by run and histogram of all lab means are included below.

Figure 1: Boxplot of lab means by run.

Figure 2: Histogram of all labs means.

## 10. Interpretation and recommended actions

- Precision: within-run and between-run precision are acceptable; sr = 0.035178, %RSD values are low.

- Trueness: T-tests show which runs differ significantly from the CRM002 value. Since all bias are significant, investigate calibration, recovery, and sample preparation.

- ANOVA result: There is a statistically significant difference between runs (p < 0.05), since significant, investigate systematic differences between analysts or dates.

## 11. Appendix - Data files

The raw data table and plots used in this report are attached as files.