| AzurESky "secure quality of coal" | RAW DATA RECORD: AS015TF-02 TWO-STAGE TOTAL MOISTURE DETERMINATION (STAGE 1 FREE MOISTURE & STAGE 2 RESIDUAL MOISTURE) | Form No: AS015TF-02 Test Method: AS015IW Cl. 2 – 6 Standard: ISO 589:2008 (Method A2) / ISO 13909-4 Effective Date: 02/2026 Revision Status: 00 |
| --- | --- | --- |

| 1. SAMPLE IDENTIFICATION & PRIMARY SPLIT (AS015IW Clause 2.1) |  |  |  |
| --- | --- | --- | --- |
| Sample ID / Barcode: |  | LIMS Job / Batch No: |  |
| Client / Source / Pit: |  | Receiving Book Entry No: |  |
| Date & Time Sampled: |  | Date & Time Received: |  |
| Nominal Top-Size (≤50mm): | [ ] Verified ≤ 50 mm | Cone & Quartering Split: | [ ] 1 Quarter [ ] 2 Opposite Quarters (≥35kg) |

| 2. STAGE 1: AIR DRYING & FREE (SURFACE) MOISTURE (AS015IW Clause 2.2) |  |  |  |
| --- | --- | --- | --- |
| Initial "As Received" Mass (m1): | ________________ kg (≥ 35 kg) | Air-Drying Method: | [ ] Ambient [ ] Fan Oven (≤40 °C) |
| Drying Pan / Tray ID(s): |  | Air-Drying Start Date/Time: |  |
| Air-Drying Stability Check: (Loss < 0.1 g between checks) | Check 1: ________ g Check 2: ________ g | Stability Check Time: Mass Loss: ________ g | [ ] Confirmed Stable (<0.1g loss) |
| Final Air-Dried Mass (m2): | ________________ kg | Air-Drying Completion Date/Time: |  |
| Surface (Free) Moisture (Mar) = [(m1 − m2) / m1] × 100 |  | Mar = ________________ % |  |

| 3. SUB-SAMPLING & CRUSHING RECORD (AS015IW Clause 3 & AS001IW) |  |  |  |
| --- | --- | --- | --- |
| Primary Crusher (25 mm) ID: |  | 25 mm Crusher Efficiency Check: | [ ] Pass (Mon shift check) |
| Riffle Splitter ID: |  | Secondary Crusher (-3 mm) ID: |  |
| -3 mm Crusher Efficiency Check: | [ ] Pass (Pre-use check) | Final Riffling (1 kg / side): | [ ] Verified (Pan A & Pan B) |

| 4. STAGE 2: OVEN DRYING & RESIDUAL MOISTURE (-3 mm) (AS015IW Clause 4.1) |  |  |  |
| --- | --- | --- | --- |
| Moisture Oven ID: |  | Oven Temp Set Point / Range: | 105 °C – 110 °C (Target: 107 °C) |
| Balance ID & Cal Expiry: |  | AS023TF Daily Verification: | [ ] In tolerance Actual: _____ °C |
| Measurement / Parameter | Unit | DUPLICATE PAN A | DUPLICATE PAN B |
| Drying Pan Identification | — |  |  |
| Mass of Empty Clean Pan Tare (mtare) | g |  |  |
| Pan Tare + Sample Before Oven Drying (mgross, initial) | g |  |  |
| Net Sample Mass Before Drying (m3) (Target: ~1000 g) | g |  |  |
| Time In Oven (105 °C – 110 °C) | hh:mm |  |  |
| Initial 3-Hour Drying Check: Mass (LIMS capture) | g |  |  |
| 30-Minute Interval Re-weighing: Mass (m3.5h) | g |  |  |
| 30-Minute Interval Mass Loss (Δm) | g |  |  |
| Stability Check (Δm ≤ 0.1 g → Stable) | status | [ ] Stable [ ] Repeat 30m | [ ] Stable [ ] Repeat 30m |
| Total Oven Drying Time (≤ 5.0 hours max) | hours |  |  |
| Final Net Stabilized Mass After Drying (m4) | g |  |  |
| Residual Moisture Loss (m3 − m4) | g |  |  |
| Residual Moisture Percentage (Mres) = [(m3 − m4) / m3] × 100 | % | Mres,A = ________ % | Mres,B = ________ % |

| 5. QUALITY CONTROL & TOTAL MOISTURE CALCULATIONS (AS015IW Cl. 5 & 6) |  |  |  |
| --- | --- | --- | --- |
| Residual Moisture Difference ¦Mres,A − Mres,B¦: | ____________ % | Repeatability Limit (2nd Stage): | ≤ 0.30 % |
| Repeatability Status (2nd Stage): | [ ] PASS [ ] REPEAT ALERT (>0.30%) | Average Residual Moisture (Mres,avg): | ____________ % |
| AzureSky LIMS Formula: TM = Mar + [(100 − Mar) × Mres,avg] / 100 LIMS Total Moisture = ________________ % |  | ISO 589 Standard Formula: TMA = Mar + [(100 − Mar) × Mres,A] / 100 = _______ % TMB = Mar + [(100 − Mar) × Mres,B] / 100 = _______ % ¦TMA − TMB¦ = _______ % (≤ 0.50% limit) |  |
| FINAL REPORTED TOTAL MOISTURE (nearest 0.1%): |  | TOTAL MOISTURE = ____________ % |  |

| 6. AUTHORIZATIONS & APPROVAL |  |
| --- | --- |
| Analyst Signature: Name: ____________________________________ Signature: ________________________________ Date: ____________ | Supervisor / Senior Analyst Approval: Name: ____________________________________ Signature: ________________________________ Date: ____________ |