# Method Validation Report

Test Method: Determination of Volatile Matter

Standard Method: ISO 562

Matrix: Coal samples

Laboratory: AzureSky

Analyst: Keitumetse Maleka

Date: 06 May to 23 June 2025

# 1. Purpose

To validate the laboratory’s method for determining volatile matter in coal in accordance with ISO 562, ensuring the method is fit for its intended use and meets ISO/IEC 17025 requirements.

# 2. Scope

This validation covers volatile matter determination for various coal types received in the laboratory, in compliance with ISO 562.

# 3. References

- ISO 562: Hard coal and coke — Determination of volatile matter

- ISO/IEC 17025:2017

- Laboratory Quality Manual

# 4. Validation Parameters & Results

| Parameter | Requirement | How to Evaluate | Acceptance Criteria | Results |
| --- | --- | --- | --- | --- |
| Repeatability | Precision under same conditions | 10 replicates on same sample | %RSD<0.3. 3% of the mean, for vol values >10% | - Repeatability:0.417% - %RSD values 0.518 and 0.484, for analysts 1 and 2; respectively. |
| Reproducibility | Precision across analysts/instruments/days | Same sample by 10 different PT participants | 0.5% or 4% of the mean, whichever is greater. | 0.327% |
| Accuracy | Comparison to CRM or known value | Test CRM with known volatile matter | Recovery 98–102% | 99.22% |
| LOD | Not typically needed for VM | Not applicable for gravimetric % methods | – | – |
| Robustness | Small variations in time/temp/sample weight | Change time from 7 min to 4 min | Difference ≤ 0.3% VM | 0.39 |
| Uncertainty | Required by ISO/IEC 17025 | Use GUM approach or Nord test Guide | k=2, 95% confidence | 0.069 |

# 5. Equipment Used

| Equipment | Model/ID | Calibration Due |
| --- | --- | --- |
| Furnace | AS01338 | March 2026 |
| Analytical Balance | AS2156 | March 2026 |
| Crucibles with lids | ISO compliant | – |

# 6. Sample Types Validated

- Bituminous coal

# 7. Validation Summary

The volatile matter determination method has demonstrated acceptable precision, accuracy, and ruggedness in accordance with ISO 562. However, the robustness has failed. This could be because the running times were too different (4 and 7 min). Measurement uncertainty has been calculated and documented.

# 8. Conclusion

The method validation fails due to failure in the statistical evaluation of robustness. Thus, the method is not suitable for its purpose. Analysis shall be repeated for 6 and 8 min; these will be compared with the 7 min standard time. Re-evaluation of the results is documented below.

# 9. Robustness re-evaluation

| t-Test: Paired Two Sample for Means |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | 7 MIN | 8 MIN |  |  |  |  |  |
| Mean | 28.506 | 28.612 |  |  |  |  |  |
| Variance | 0.0223175 | 0.00952 |  |  |  |  |  |
| Observations | 5 | 5 |  |  |  |  |  |
| Pearson Correlation | 0.273392804 |  |  |  |  |  |  |
| Hypothesized Mean Difference | 0 |  |  |  |  |  |  |
| df | 4 |  |  |  |  |  |  |
| t Stat | -1.534219159 |  |  |  |  |  |  |
| P(T<=t) one-tail | 0.099880486 |  |  |  |  |  |  |
| t Critical one-tail | 2.131846786 |  |  |  |  |  |  |
| P(T<=t) two-tail | 0.199760972 |  |  |  |  |  |  |
| t Critical two-tail | 2.776445105 |  |  |  |  |  |  |
| Evaluation |  |  |  |  |  |  |  |
| Ho: | There is no significant difference b/w the means of 7- and 8-min results |  |  |  |  |  |  |
| Ha: | There is a significant difference b/w the means of 7- and 8-min results |  |  |  |  |  |  |
| tcalculated: | 0.099880486 |  |  |  |  |  |  |
| p-value (2 tail) : | 0.199760972 |  |  |  |  |  |  |
| tcritical | 2.776445105 |  |  |  |  |  |  |
| Since the p-value is greater than 0.05, and the tcalculated is less than tcritical. Accept the null hypothesis. |  |  |  |  |  |  |  |

# Conclusion: There is no significant difference b/w the 7 min and 8 min.  This implies that for volatile analysis, samples left in the furnace for 7-8 min would produce results that are close to each other.

| LIMS NO | 6 MIN | 7 MIN | 8 MIN | LIMS NO |
| --- | --- | --- | --- | --- |
| 1268 | 28.57 | 28.65 | 28.76 | 1273 |
| 1269 | 28.56 | 28.43 | 28.56 | 1274 |
| 1270 | 28.55 | 28.345 | 28.66 | 1275 |
| 1271 | 28.26 | 28.425 | 28.52 | 1276 |
| 1272 | 28.31 | 28.68 | 28.56 | 1277 |

| t-Test: Paired Two Sample for Means |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | 7 MIN | 6 MIN |  |  |  |  |
| Mean | 28.506 | 28.45 |  |  |  |  |
| Variance | 0.0223175 | 0.02305 |  |  |  |  |
| Observations | 5 | 5 |  |  |  |  |
| Pearson Correlation | -0.178014057 |  |  |  |  |  |
| Hypothesized Mean Difference | 0 |  |  |  |  |  |
| df | 4 |  |  |  |  |  |
| t Stat | 0.541663593 |  |  |  |  |  |
| P(T<=t) one-tail | 0.308405515 |  |  |  |  |  |
| t Critical one-tail | 2.131846786 |  |  |  |  |  |
| P(T<=t) two-tail | 0.61681103 |  |  |  |  |  |
| t Critical two-tail | 2.776445105 |  |  |  |  |  |
| Evaluation |  |  |  |  |  |  |
| Ho: | There is no significant difference b/w the means of 6- and 7-min results |  |  |  |  |  |
| Ha: | There is a significant difference b/w the means of 6- and 7-min results |  |  |  |  |  |
| tcalculated: | 0.541663593 |  |  |  |  |  |
| p-value (2 tail) : | 0.61681103 |  |  |  |  |  |
| tcritical | 2.776445105 |  |  |  |  |  |
| Since the p-value is greater than 0.05, and the tcalculated is less than tcritical. Accept the null hypothesis. |  |  |  |  |  |  |
| # Conclusion: There is no significant difference b/w the 7 min and 6 min. This implies that for volatile analysis, samples left in the furnace for 7-6min would produce results that are close to each other. |  |  |  |  |  |  |

# 10. Glossary

Terms to know and why:

- Robustness: The method’s ability to remain unaffected by small, deliberate changes in method parameters- such as: Holding time, furnace temperature, Sample mass, lid coverage or crucible type.

Why is it important: In real laboratories slight deviations do happen, instruments may drift slightly, or timing may vary b/w runs. You want to know if the method can tolerate those changes without producing significant errors. This means that results are dependent on operator precision or timing, which introduces systematic errors.