## RAW DATA
LIMS no | 7 Min |  | Average (7 min) | LIMS no | 4 Min |  | Average (4 min)
1208 | 28.67 | 28.63 | 28.65 | 1258 | 28.03 | 28.26 | 28.145
1209 | 28.38 | 28.48 | 28.43 | 1259 | 28.13 | 28.18 | 28.155
1210 | 28.3 | 28.39 | 28.345 | 1260 | 28.01 | 28.03 | 28.02
1211 | 28.35 | 28.5 | 28.425 | 1261 | 27.97 | 28.04 | 28.005
1212 | 28.64 | 28.72 | 28.68 | 1262 | 28.01 | 28.07 | 28.04
1213 | 28.34 | 28.39 | 28.365 | 1263 | 27.96 | 27.96 | 27.96
1214 | 28.26 | 28.23 | 28.245 | 1264 | 27.98 | 28 | 27.99
1215 | 28.27 | 28.28 | 28.275 | 1265 | 27.8 | 27.88 | 27.84
1216 | 28.24 | 28.38 | 28.31 | 1266 | 27.85 | 28.06 | 27.955
1217 | 28.36 | 28.57 | 28.465 | 1267 | 28.13 | 28.14 | 28.135

## DATA SUMMARY
Index | Average (4 min) | Average (7 min) | Diff(7-4)
1 | 28.145 | 28.65 | 0.505
2 | 28.155 | 28.43 | 0.275
3 | 28.02 | 28.345 | 0.325
4 | 28.005 | 28.425 | 0.42
5 | 28.04 | 28.68 | 0.64
6 | 27.96 | 28.365 | 0.405
7 | 27.99 | 28.245 | 0.255
8 | 27.84 | 28.275 | 0.435
9 | 27.955 | 28.31 | 0.355
10 | 28.135 | 28.465 | 0.33
| 28.0245 | 28.419 | 0.3945

## LINEARITY
Since we're comparing two conditions (4 min vs 7 min) with the same sample set, linearity isn't directly applicable unless these are collected over a concentration or time gradient. In this case, you are more validating method robustness/ruggedness.

## PRECISION & ACCURACY
Repeatability
Air-dry basis
Same normal conditions |  |  |  |  | Assigned value | 28.64
| Analyst 1 |  | Analyst 2 |  | Obtained Average | 28.419
LIMS no | Average | LIMS no | Average |  | Difference | 0.221 | ACCURACY
1208 | 28.65 | 143 | 28.85 |  | Acceptance | ≤0.5%
1209 | 28.43 | 144 | 28.87 |  | Recovery | 99.2284
1210 | 28.345 | 145 | 29.21
1211 | 28.425 | 146 | 28.77 |  | Reproducibility limit is 0.5% or 4% of the mean result, whichever is greater. Thus, the method passes the accuracy test
1212 | 28.68 | 147 | 28.78
1213 | 28.365 | 148 | 28.7
1214 | 28.245 | 149 | 28.83
1215 | 28.275 | 150 | 28.79
1216 | 28.31 | 151 | 28.79
1217 | 28.465 | 152 | 28.77
Avergae | 28.419 |  | 28.836
Diff |  | 0.417 |  | The RSD values obtained are attributed to very high precision and low variability in the data.
SD | 0.14721 |  | 0.13978
%RSD | 0.518 |  | 0.48473
Assigned value | 28.64
Reproducibility
Mean | 28.419 | 28.746
SD | 0.14721 | 0.4596
%RSD | 0.518 | 1.5988
Reproducibility | 0.327

## RUGGEDNESS AND ROBUSTNESS
[11 empty] | LIMS no | 7 Min |  | Average (7 min) | LIMS no | 4 Min |  | Average (4 min)
Robustness: Here we assess method sensitivity to small changes (e.g., duration of 4 min vs 7 min) | [10 empty] | 1208 | 28.67 | 28.63 | 28.65 | 1258 | 28.03 | 28.26 | 28.145
Ruggedness: Assess variations across labs, analysts, equipment etc. | [7 empty] | t-Test: Paired Two Sample for Means |  |  | 1209 | 28.38 | 28.48 | 28.43 | 1259 | 28.13 | 28.18 | 28.155
[11 empty] | 1210 | 28.3 | 28.39 | 28.345 | 1260 | 28.01 | 28.03 | 28.02
[9 empty] | Average (7 min) | Average (4 min) | 1211 | 28.35 | 28.5 | 28.425 | 1261 | 27.97 | 28.04 | 28.005
Robustness: | [7 empty] | Mean | 28.419 | 28.0245 | 1212 | 28.64 | 28.72 | 28.68 | 1262 | 28.01 | 28.07 | 28.04
Looking at the data summary, there is a consistent shift in the results; with 7 min values being higher on all occasions. The consistent shift in means indicate sensitivity to this parameter. | [7 empty] | Variance | 0.021671 | 0.0098469 | 1213 | 28.34 | 28.39 | 28.365 | 1263 | 27.96 | 27.96 | 27.96
[8 empty] | Observations | 10 | 10 | 1214 | 28.26 | 28.23 | 28.245 | 1264 | 27.98 | 28 | 27.99
[8 empty] | Pearson Correlation | 0.62557 |  | 1215 | 28.27 | 28.28 | 28.275 | 1265 | 27.8 | 27.88 | 27.84
tcalculated | 10.8413 | [6 empty] | Hypothesized Mean Difference | 0 |  | 1216 | 28.24 | 28.38 | 28.31 | 1266 | 27.85 | 28.06 | 27.955
p-value | 1.8197e-06 | [6 empty] | df | 9 |  | 1217 | 28.36 | 28.57 | 28.465 | 1267 | 28.13 | 28.14 | 28.135
tcritical | 2.2622 | [6 empty] | t Stat | 10.8413
[8 empty] | P(T<=t) one-tail | 9.0983e-07
There is a significant difference b/w the results. Thus, the 4 min vs 7 min avergaes are not interchangeable. Duration will affect the final volatile reading significantly. | [7 empty] | t Critical one-tail | 1.8331 |  |  | t-Test: Paired Two Sample for Means
[8 empty] | P(T<=t) two-tail | 1.8197e-06
[8 empty] | t Critical two-tail | 2.2622 |  |  |  | Average (7 min) | Average (4 min)
[12 empty] | Mean | 28.419 | 28.0245
[12 empty] | Variance | 0.021671 | 0.0098469
Ruggedness | [11 empty] | Observations | 10 | 10
ISO 562 Volatile Matter Determination | [11 empty] | Pearson Correlation | 0.62557
LIMS no | Azuresky Average (7 min) | 10 Random points from PT results (different labs) for the month of may |  | LAB ID | Differences | t-Test: Paired Two Sample for Means | [5 empty] | Hypothesized Mean Difference | 0
[12 empty] | df | 9
1208 | 28.65 | 28.98 |  | 16 | 0.33 |  | Variable 1 | Variable 2 |  |  |  | t Stat | 10.8413
1209 | 28.43 | 29.7 |  | 45 | 1.27 | Mean | 28.419 | 28.746 |  |  |  | P(T<=t) one-tail | 9.0983e-07
1210 | 28.345 | 28.9 |  | 21 | 0.555 | Variance | 0.021671 | 0.21127 |  |  |  | t Critical one-tail | 1.8331
1211 | 28.425 | 28.92 |  | 26 | 0.495 | Observations | 10 | 10 |  |  |  | P(T<=t) two-tail | 1.8197e-06
1212 | 28.68 | 28.2 |  | 39 | -0.48 | Pearson Correlation | -0.0032678 |  |  |  |  | t Critical two-tail | 2.2622
1213 | 28.365 | 28.49 |  | 41 | 0.125 | Hypothesized Mean Difference | 0
1214 | 28.245 | 28.1 |  | 44 | -0.145 | df | 9
1215 | 28.275 | 28.92 |  | 26 | 0.645 | t Stat | -2.1405
1216 | 28.31 | 28.77 |  | 42 | 0.46 | P(T<=t) one-tail | 0.030487
1217 | 28.465 | 28.48 |  | 22 | 0.015 | t Critical one-tail | 1.8331
[6 empty] | P(T<=t) two-tail | 0.060973
Mean | 28.419 | 28.746 | #DIV/0! |  |  | t Critical two-tail | 2.2622
Standard deviation | 0.14721 | 0.45964 | #DIV/0!
Evaluation
[6 empty] | Ho: | There is no significant difference b/w the means of azuresky and PT participants
[6 empty] | Ha: | There is a significant difference b/w the means of azuresky and PT participants
[6 empty] | tcalculated: | -2.1405
[6 empty] | p-value (2 tail) : | 0.060973
[6 empty] | tcritical | 2.2622
Since the p-value is greater than 0.05, and the tcalculated is less than tcritical. Accept the null hypothesis.
Conclusion: There is no statistically significant difference b/w azuresky results and PT participants at CL of 95%. Although there is a numerical difference b/w the two means, it is not large enough to be considered statistically significant given the sample size and variability.

## ROBUSTNESS RE-EVALUATION
LIMS NO | 6 MIN | 7 MIN | 8 MIN | LIMS NO | diff(7-6) | diff(8-7)
1268 | 28.57 | 28.65 | 28.76 | 1273 | 0.08 | 0.11
1269 | 28.56 | 28.43 | 28.56 | 1274 | -0.13 | 0.13
1270 | 28.55 | 28.345 | 28.66 | 1275 | -0.205 | 0.315
1271 | 28.26 | 28.425 | 28.52 | 1276 | 0.165 | 0.095
1272 | 28.31 | 28.68 | 28.56 | 1277 | 0.37 | -0.12
t-Test: Paired Two Sample for Means
t-Test: Paired Two Sample for Means
[8 empty] | 7 MIN | 6 MIN
| 7 MIN | 8 MIN |  |  |  |  | Mean | 28.506 | 28.45
Mean | 28.506 | 28.612 |  |  |  |  | Variance | 0.022317 | 0.02305
Variance | 0.022317 | 0.00952 |  |  |  |  | Observations | 5 | 5
Observations | 5 | 5 |  |  |  |  | Pearson Correlation | -0.17801
Pearson Correlation | 0.27339 | [5 empty] | Hypothesized Mean Difference | 0
Hypothesized Mean Difference | 0 | [5 empty] | df | 4
df | 4 | [5 empty] | t Stat | 0.54166
t Stat | -1.5342 | [5 empty] | P(T<=t) one-tail | 0.30841
P(T<=t) one-tail | 0.09988 | [5 empty] | t Critical one-tail | 2.1318
t Critical one-tail | 2.1318 | [5 empty] | P(T<=t) two-tail | 0.61681
P(T<=t) two-tail | 0.19976 | [5 empty] | t Critical two-tail | 2.7764
t Critical two-tail | 2.7764
Evaluation | [6 empty] | Evaluation
Ho: | There is no significant difference b/w the means of 7 and 8 min results | [5 empty] | Ho: | There is no significant difference b/w the means of 6 and 7 min results
Ha: | There is a significant difference b/w the means of 7 and 8 min results | [5 empty] | Ha: | There is a significant difference b/w the means of 6 and 7 min results
tcalculated: | 0.09988 | [5 empty] | tcalculated: | 0.54166
p-value (2 tail) : | 0.19976 | [5 empty] | p-value (2 tail) : | 0.61681
tcritical | 2.7764 | [5 empty] | tcritical | 2.7764
Since the p-value is greater than 0.05, and the tcalculated is less than tcritical. Accept the null hypothesis. | [6 empty] | Since the p-value is greater than 0.05, and the tcalculated is less than tcritical. Accept the null hypothesis.
Conclusion: There is no significant difference b/w the 7 min and 8 min. This implies that for volatile analysis, samples left in the furnace for 7-8 min would produce results that are close to each other. | [6 empty] | Conclusion: There is no significant difference b/w the 7 min and 6 min. This implies that for volatile analysis, samples left in the furnace for 7-6min would produce results that are close to each other.

## UoM
7 minutes
| 1 | 2 | Anova: Single Factor
| 28.67 | 28.63
| 28.38 | 28.48 | SUMMARY
| 28.3 | 28.39 | Groups | Count | Sum | Average | Variance
| 28.35 | 28.5 | 1 | 10 | 283.81 | 28.381 | 0.02301
| 28.64 | 28.72 | 2 | 10 | 284.57 | 28.457 | 0.023557
| 28.34 | 28.39
| 28.26 | 28.23
| 28.27 | 28.28 | ANOVA
| 28.24 | 28.38 | Source of Variation | SS | df | MS | F | P-value | F crit
| 28.36 | 28.57 | Between Groups | 0.02888 | 1 | 0.02888 | 1.2404 | 0.28005 | 4.4139
Mean | 28.419 |  | Within Groups | 0.4191 | 18 | 0.023283
SD | 0.15355
Sswithin | 0.4191 |  | Total | 0.44798 | 19
Ssbetween | 0.02888
Check value | 1.3091
Since check value> ssbetween, homogeneity test passes.
Standard Uncertainty | 0.034335
Expanded uncertainty (k=2) | 0.06867