Contents

1. Scope and Technical Objective	2

2. Process Train and Equipment	2

3. Evaluation of Performance Characteristics	3

4. Statistical Methodology and Outlier Detection	5

5. Uncertainty of Measurement (UoM) Assessment	5

6. Mandatory Operational Controls	6

7. Validation Maintenance	7

8. Maintenance of Uncertainty Budget	15

9. Overall Summary	16

## 1. Scope and Technical Objective

The objective of this validation is to quantify the performance characteristics of the physical mass reduction and homogenization process for coal. This ensures that the preparation variance does not negatively impact the target standard deviation (Sigma pt) of Proficiency Testing (PT) schemes or the certified values of Reference Materials (CRMs).

## 2. Process Train and Equipment

- Mass Reduction (Stage Crushing): Platform scales, 25mm, 10mm, and 3mm/4.75mm crushers.

- Homogenization: 210L rotating drum.

- Final Milling: Pulverizer for reduction to -212mm.

## 3. Evaluation of Performance Characteristics

| Characteristic | Operational Status | Technical Justification |
| --- | --- | --- |
| Precision (Repeatability) | Included | Assessed via duplicate preparation streams. It identifies the variance introduced by stage-crushing and splitting (Gy’s Sampling Theory). |
| Bias / Accuracy | Included | Evaluated using Student’s T test and F test on two separate legs of the same sample. |
| Homogeneity | Included | Mandatory for PT/CRM production. Assessed via a 10-bottle random sampling plan from the 210L drum to ensure between-bottle uniformity. |
| Stability X | Included | Evaluation of coal parameters over time in the storage area to verify shelf-life for CRMs and PT transit times. |
| Robustness | Included | Evaluated through equipment efficiency checks. The 3mm and 4.75mm crusher efficiency must be verified every time before use to prevent PSD failure. |
| Uncertainty (UoM) | Included | A "Top-Down" assessment combining preparation, homogeneity, and analytical variance into a single expanded uncertainty budget (k=2). |
| Linearity | Omitted | Physical preparation is a mass-reduction process; it does not involve concentration-response curves or instrument signals. |
| LOD / LOQ | Omitted | Applies only to analytical sensitivity for trace elements, not the physical reduction of bulk coal. |
| Selectivity | Omitted | Concerns the ability of a method to distinguish between analytes, which is irrelevant to physical grinding. |

## 4. Statistical Methodology and Outlier Detection

The laboratory employs robust statistical methods to process validation and PT data:

- Algorithm A (ISO 13528): Used to calculate the robust mean and standard deviation, minimizing the influence of extreme values without simple deletion.

- Z-Scores: Performance in PT rounds is evaluated using Z-scores. Any value where Z > 3 is classified as an outlier and triggers a technical investigation in the Risk Register (Form ASRI 8.5-001).

- Homogeneity Assessment: The between-bottle standard deviation must meet the criterion.

## 5. Uncertainty of Measurement (UoM) Assessment

The preparation uncertainty (u_prep) is a critical component of the laboratory's technical signatory requirements. The expanded uncertainty (U) is calculated as:

U = k sqrt{u_{prep}^2 + u_{hom}^2 + u_{anal}^2}

- u_prep: Standard deviation of duplicate preparation replicates.

- u_hom: Between-bottle variance from the homogeneity study.

- u_anal: The laboratory’s established repeatability for the specific coal parameter (e.g., Ash, Sulphur).

## 6. Mandatory Operational Controls

To maintain the validated state of method AS001IW, the following controls are implemented:

- Crusher Efficiency Check: Moved from a weekly schedule to a "Per-Use" requirement for both 3mm and 4.75mm aperture crushers.

- Homogenization Duration: A minimum 4-hour drum cycle is required for bulk lots (approx. 200kg) to ensure particle distribution across the drum volume.

- Drying Temperature: Air-drying must be strictly maintained at 40 °C to prevent the oxidation of volatile matter.

- Hold/Release Points: If CRM ID stickers or specific labels are unavailable, interim controls must be applied to the batch to prevent the misassignment of materials during packaging.

## 7. Validation Maintenance

This validation is reviewed:

- Annually during Management Review.

- Upon significant maintenance or change in preparation equipment (e.g., replacement of crusher wear plates).

PERFORMANCE CHARACTERISTICS

- Homogeneity

- Purpose: To ensure the 4-hour homogenization in the 210 L drum eliminates variance.

- Procedure:  Collect 10 random bottles whilst packaging a batch.

- Analyze each bottle in duplicate for Ash content (the primary indicator of homogeneity).

- Statistical Evaluation:  Calculate the Mean and Standard Deviation.

- Perform Anova using excel.

- Acceptance Criteria: SSbetween<check value

- Repeatability

Procedure: A bulk sample was selected & split at the initial stage of the preparation using the cone & quarter method. They are namely leg one & leg two. Each of the legs was processed following the preparation procedure AS001IW. 10 Random bottles were selected, which were used to analyse Inherent moisture & Ash for evaluation.

Evaluation: Use the differences between the duplicate pairs to calculate the preparation standard deviation.

| ` |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Dry basis |  |  |  |  |  |
| # Sample | %R1 | %R2 | %Sample av (Xt) | Range (Wt) | Range sqd |  |
| 1825 | 22,80 | 22,55 | 22,67 | 0,2561 | 0,0656 |  |
| 1826 | 22,55 | 22,73 | 22,64 | 0,1893 | 0,0358 |  |
| 1827 | 22,92 | 22,78 | 22,85 | 0,1414 | 0,0200 |  |
| 1828 | 22,01 | 22,08 | 22,04 | 0,0704 | 0,0050 |  |
| 1829 | 23,22 | 23,22 | 23,22 | 0,0015 | 0,0000 |  |
| 1830 | 21,86 | 22,22 | 22,04 | 0,3504 | 0,1228 |  |
| 1831 | 22,31 | 22,76 | 22,53 | 0,4512 | 0,2036 |  |
| 1832 | 22,73 | 23,06 | 22,90 | 0,3327 | 0,1107 |  |
| 1833 | 22,80 | 22,84 | 22,82 | 0,0380 | 0,0014 |  |
| 1834 | 22,57 | 22,51 | 22,54 | 0,0587 | 0,0034 |  |
| Overall Average |  |  | 22,63 |  |  |  |
| Standard Deviation |  |  | 0,38 |  |  |  |
| SSwithin |  |  | 2,667 |  |  |  |
| SSbetween |  |  | 0,0475 |  |  |  |
| Check value |  |  | 1.481 |  |  |  |
| Conclusion: SSbetween< Check value. Thus, homogeneity is sufficient. |  |  |  |  |  |  |

“LEG ONE”

| ` |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | Dry basis |  |  |  |  |  |
| # Sample | %R1 | %R2 | %Sample av (Xt) | Range (Wt) | Range sqd |  |
| 1 | 22,66 | 22,81 | 22,73 | 0,1505141 | 0,0227 |  |
| 2 | 22,67 | 22,73 | 22,70 | 0,0528541 | 0,0028 |  |
| 3 | 22,85 | 22,77 | 22,81 | 0,0772957 | 0,0060 |  |
| 4 | 22,85 | 22,72 | 22,78 | 0,1282999 | 0,0165 |  |
| 5 | 22,56 | 22,61 | 22,58 | 0,0562482 | 0,0032 |  |
| 6 | 22,79 | 22,89 | 22,84 | 0,1007301 | 0,0101 |  |
| 7 | 22,74 | 22,82 | 22,78 | 0,0763109 | 0,0058 |  |
| 8 | 22,69 | 23,14 | 22,91 | 0,4491311 | 0,2017 |  |
| 9 | 23,27 | 23,15 | 23,21 | 0,1283765 | 0,0165 |  |
| 10 | 23,18 | 23,39 | 23,29 | 0,209926 | 0,0441 |  |
| Overall Average |  |  | 22,86 |  |  |  |
| Standard Deviation |  |  | 0,23 |  |  |  |
| SSwithin |  |  | 1,016 |  |  |  |
| SSbetween |  |  | 0,0290 |  |  |  |
| Õbetween |  |  | 0,7025 |  |  |  |
| Check value |  |  | 1,235 |  |  |  |
| Conclusion :Õbetween< Check value. Thus, homogeneity is sufficient. |  |  |  |  |  |  |

“LEG TWO”

The decision rule has been met; therefore, the sample has been sufficiently homogenized

Evaluation

The standard deviation of duplicates (the "Difference between pairs" method) is used to determine the precision of the preparation process. The formula is:

õprep = √(Mean Difference)22n

= √(22.86-22.63)22(20)

= 0.115

‘To detect if the method does not take up too much of the error budget, the standard deviation is compared to the most check value from

‘ a PT report, Novemeber PT B.(1.481)

Therefore: Preparation ≤ 0.3 x õpt

The standard deviation of the preparation process (u_prep = 0.115) was evaluated against the PT target standard deviation (sigma_pt = 1.481). The ratio u_prep ≤ sigma_pt x 0.3, confirming that the variation introduced during the crushing, splitting, and milling stages is statistically insignificant relative to the total allowable error of the proficiency testing scheme. Consequently, the preparation process is validated as fit for purpose, and its contribution to the overall measurement uncertainty is considered negligible.

Students’ t-test & F test

| Test | Required Value | Constant (n=10) |
| --- | --- | --- |
| t-test | t_crit (0.05, 18 df, 2-tail) | 2.101 |
| F-test | F_crit (0.05, 9 df, 9 df) | 3.179 |

Rule: If  |t_calc| > 2.101, the result is Failed (Significant Bias detected in splitting).

Rule: If |t_calc| < 2.101, the result is Passed (No significant bias)

Calculation(tcrit)):

Leg 1: Mean is 22.63; SD is 0.38

Leg 2: Mean is 22.86; SD is 0.23

Standard Error= s12n1+s22n2+

Where s is the standard deviation of the lot & n is the sample size

=0.144410+0.052910

=0.1405

tcritical= x1-x2SE

=22,63-22,830,1405

= 1.42

Decision Rule:

Compare t_calc to the t-critical value (2.101).

- Result: 1.42 < 2.101

- Conclusion: Passed. There is no statistically significant bias between the two legs. The splitting method (Cone and Quartering) is validated as representative.

F test

‘F= s2maxs2min

‘=0.3820.232

= 2.73

Decision Rule:

Compare F_calc to the F-critical value (3.179).

- Result: 2.73 < 3.179

- Conclusion: Passed. The precision is consistent across both legs. The preparation process is stable and repeatable

Homogeneity Test

The two used legs have been tested for homogeneity, wherein 10 random bottles have been selected for analysis during packaging as per the procedure AS001IW. The decision rules in a homogeneity test are that  Conclusion : Õbetween< Check value. Thus, homogeneity is sufficient.

Uncertainty of Measurement (UoM) Assessment: Method AS001IW

1. Calculation Model

As per the validated framework, the expanded uncertainty (U) is derived from a "Top-Down" assessment. This model combines the variance introduced by the physical preparation, the homogenization of the bulk lot, and the analytical repeatability of the laboratory.

The expanded uncertainty is calculated using the following formula:

U = k √u_prep + u_hom2 + u_anal2

- k: The coverage factor, set at 2 for a 95% confidence level.

- u_prep: Preparation uncertainty.

- u_{hom: Homogeneity (between-bottle) uncertainty.

- u_anal: Analytical repeatability.

2. Uncertainty Components (Values from Validation Data)

| Component | Description | Established Value | Technical Source |
| --- | --- | --- | --- |
| Preparation (u_prep) | Standard deviation of duplicate preparation replicates (Leg 1 vs. Leg 2). | 0.115 | Difference between pairs method. |
| Homogeneity (u_hom) | Between-bottle variance assessed via 10-bottle random sampling from the 210L drum. | 0.0290 | Leg 1 homogeneity SS_between. |
| Analytical (u_anal) | Laboratory repeatability for the specific coal parameter (using Ash as the primary indicator). | 0.23 | Precision established for Leg 2 replicates. |

3. Calculation of Expanded Uncertainty (U)

To determine the combined standard uncertainty (u_c):

- Square the components:

  - u_prep2 = (0.115)2 = 0.0132

  - u_hom2 = (0.0290)2 = 0.0008

  - u_anal2 = (0.23)2 = 0.0529

- Sum of squares: 0.0132 + 0.0008 + 0.0529 = 0.0669

- Combined Standard Uncertainty (u_c): √0.0669 = 0.2587

- Expanded Uncertainty (U):0.2587 X 2 = 0.517

Final Expanded Uncertainty (k=2) 0.52

## 8. Maintenance of Uncertainty Budget

This budget must be reviewed annually or upon significant maintenance, such as the replacement of crusher wear plates. Any technical investigation in the Risk Register (Form ASRI 8.5-001) resulting from a Z-score > 3 in PT rounds may require a re-evaluation of these components.

Conclusion:

The precision of the sample preparation method AS001IW was evaluated using a duplicate-leg study (n=10 per leg). Statistical analysis via the F-test confirmed consistency in preparation variance between runs (F_calc = 2.73 < F_crit = 3.18). The student’s t-test indicated no significant bias between the splitting legs (t_calc = 1.42 < t_crit = 2.10), validating the Cone and Quartering technique as representative. The resulting preparation uncertainty (u_prep = 0.115) is within the acceptable limits for Ash determination (ISO 1171), confirming the method is fit for purpose for CRM production and PT schemes.

NOTE: Though the method is deemed fit for this performance characteristic, it should be noted that cone & quartering heavily depends on human involvement, which is then prone to many human errors. It is therefore documented that when a sample undergoes cone & quartering, the two opposites that remain at that stage will only be referred to as “reference sample to the original” once it undergoes preparation, either internally analysed or sent for PT distribution, and the results are reproducible to the original sample.

## 9. Overall Summary

The expanded uncertainty (U) for reparation is established at 0.52 (k=2). This budget integrates variances from physical mass reduction, homogenization, and analytical repeatability. Statistical evaluation confirms that the preparation uncertainty (uprep = 0.115) contributes negligibly to the total error, as it remains well below the 0.3 ×σpt threshold established from proficiency testing. Method AS001IW is therefore validated as Fit for Purpose for CRM production and PT schemes. This is supported by F-test results (2.73 < 3.18) and t-test results (1.42 < 2.10), which confirm consistent precision and a lack of significant bias in the manual splitting process. To maintain this validated status, the laboratory must adhere to mandatory "Per-Use" crusher efficiency checks and strictly controlled air-drying temperatures of 40ºC.