## MODEL ANSWERS AND MARKING RUBRIC (TOTAL: 20 POINTS)

| Q# | Question Focus | Model Answer & Mandatory Technical Criteria | Points |
| --- | --- | --- | --- |
| 1 | Method Selection & Principle (AS015IW Cl. 1, ISO 589 Cl. 4) | Two methods: Single-Stage and Two-Stage thermal drying (1 pt). Requirement for Two-Stage: Required when coal is visibly wet / excessively moist, or when particle size/mass makes single-stage drying impractical or prone to moisture loss/choking during crushing (1 pt). | 2 pts |
| 2 | Sample Mass & Top-Size (AS015IW Cl. 2.1, ISO 13909-4 Table 1, AS001IW Cl. 4.1) | Minimum mass for 50 mm top-size: 35 kg (1 pt). Maximum nominal top-size accepted by lab: Shall not be > 50 mm (1 pt). | 2 pt |
| 3 | Cone & Quartering (AS015IW Cl. 2.1) | Form coal into cone, flatten top, divide into 4 equal quarters, select one quarter (1 pt). If the selected quarter mass < 35 kg, the analyst must combine two opposite quarters to ensure representativeness (1 pt). | 2 pts |
| 4 | Stage 1 Air-Drying Conditions (AS015IW Cl. 2.2, ISO 13909-4 Cl. 9, ISO 589 Cl. 6.1.1) | Air dry at ambient temperature or in fan-forced oven not exceeding 40 °C (1 pt). Rationale: Temperatures > 40 °C promote coal oxidation (especially in lower-rank coals), leading to mass gain and severe analytical bias (1 pt). | 2 pt |
| *5 | CRITICAL Procedural Flow & Preparation (AS015IW Cl. 2.2, 3.1, 3.2) | Mandatory 5 steps (1 pt each - ALL must be correct): 1. Spread 35 kg sample thinly on clean drying pans until "visibly dry" (≤ 40 °C). 2. Stage 1 stability check: air drying complete when mass loss is < 0.1 g. 3. Primary crush through 25 mm crusher and reduce mass using Riffle Splitter. 4. Secondary crush: crush to -10 mm for Single-Stage, or -3 mm for Two-Stage. 5. Final riffle division to 1 kg per side, tare pans on balance, transfer into duplicate Pan A and Pan B. | *5 pts (Zero tolerance) |
| *6 | CRITICAL Oven Temperature & Stability (AS015IW Cl. 4.1, ISO 589 Cl. 4) | Temperature range: 105 °C to 110 °C (1 pt). Drying intervals & stability: Initial dry for 3 hours; first weighing on LIMS, return for 30 minutes. Stable when mass loss between 30-min intervals is ≤ 0.1 g (1 pt). | *2 pt (Zero tolerance) |
| 7 | Oven Failure Limit (AS015IW Cl. 4.1.5) | Limit: If sample does not stabilize within 5 hours (0.5 pt). Action: Test marked as "Failed" in LIMS and must be repeated completely (0.5 pt). | 1 pt |
| 8 | Single-Stage vs Two-Stage Oven Drying (AS015IW Cl. 4.2) | In Single-Stage (-10 mm), the sample remains in the oven for the full 5-hour duration without the 3-hour initial lapse check, after which final stabilized mass is captured on LIMS (1 pt). | 1 pt |
| 9 | Equipment Verification (AS023TF, AS001IW Cl. 4.1) | Oven verification: Calibrated thermocouple reading (average + correction factor), indicator set point checked against tolerance 105–110 °C (nominal ≤ 107 °C) (1 pt). Crusher efficiency: 25 mm & 10 mm every Monday before shift; 3 mm & 4.75 mm on day of use before use (1 pt). | 2 pt |
| 10 | Free Moisture Calculation (AS015IW Cl. 5.1) | Formula: Mar = [(m1 − m2) / m1] × 100 (1 pt). Calculation: [(35.00 − 32.55) / 35.00] × 100 = (2.45 / 35.00) × 100 = 7.00% (1 pt). | 2 pts |
| 11 | Precision & Repeatability Alert (AS015IW Cl. 6, ISO 589 Table 1) | Repeatability check: ¦4.58% − 4.20%¦ = 0.38%. Since 0.38% > 0.30%, it DOES NOT comply (1 pt). LIMS Alert & Action: LIMS triggers a "repeat" alert; the sample test is invalidated and must be repeated (1 pt). | 2 pts |
| 12 | LIMS vs ISO 589 Alignment (AS015IW Cl. 6 Note) | Difference: LIMS averages the duplicate residual moistures first and inserts into TM formula: TM = Mar + [(100 − Mar) × Mres,avg]/100. ISO 589 calculates TM of Pan A and Pan B individually and averages them at the end. Mathematically both yield identical results and both are approved at AzureSky (2 pt). | 2 pt |

ASSESSMENT DECISION RULES:

- If Candidate Score ≥ 16 (80%) AND Questions *5 and *6 are 100% correct → COMPETENT.

- If Question *5 OR Question *6 is incorrect (even if score ≥ 16) → NOT YET COMPETENT (Immediate assessment stoppage).

- If Candidate Score < 16 → NOT YET COMPETENT.