Table of Contents

1. Scope and Principle	2

2. Sample Preparation and Division	2

3. Sub-Sampling and Crushing (ISO 13909-4)	3

4. Oven Drying (Residual Moisture)	4

5. Calculations	5

6. Precision Limits (ISO 589)	6

7. Equipment Verification and Quality Control	7

Revision Table	7

## 1. Scope and Principle

This procedure describes the determination of total moisture in hard coal using thermal drying in air. The procedure specifies an AzureSky In-House Validated Method based on the principles of ISO 589:2008 and ISO 13909-4:2025, operating under AzureSky's ISO/IEC 17025 management system.

Moisture is determined gravimetrically by the loss in mass of the sample when heated in air at 105 ºC to 110 °C under controlled conditions until constant mass is attained. This method is applicable to hard coals not susceptible to significant thermal oxidation.

Governing Standards & References: ISO 589:2008 (Hard coal — Determination of total moisture), ISO 13909-4:2025 (Coal — Preparation of test samples), ISO 18283, AS001IW, AS023TF.

Method Selection & Scope Criteria:
• Two-Stage Determination (Clause 2, 3 & 4.1): Applied to unreduced bulk samples, visibly wet coals, or excessively moist consignments where direct mechanical reduction would cause moisture loss by evaporation, choking of crushers, or screen blinding. Where free water separates from the coal in the sample container, the whole sample and its container shall be air-dried together in accordance with ISO 13909-4:2025 and ISO 589 Clause 7.5, and the mass loss accounted for.
• Single-Stage Determination (Clause 4.2): Applied to free-flowing, drier coals received or prepared directly at nominal top-size -10 mm.

## 2. Sample Preparation and Division

2.1. Primary Split (Bulk Sample)

- Sample Acceptance: The nominal top-size of incoming coal accepted by the laboratory shall not exceed 50 mm (AS001IW Cl. 4.1). For a 50 mm top-size bulk sample, the minimum required mass is 35 kg (ISO 13909-4 Table 1).

- Cone and Quartering: Form the coal into a cone, flatten the top uniformly, and divide into four quarters.

3. Division & Representativeness: Select one quarter. If the mass of the chosen single quarter is less than the mandatory 35 kg minimum sample mass, combine two opposite quarters to ensure statistical representativeness.

4. Initial Bulk Mass (m_1): Immediately weigh the divided bulk sample on the calibrated platform balance and record the mass on LIMS and raw data sheet AS015TF-02. This is the initial "As Received" bulk mass (m_1, in kg).

2.2. Stage 1: Air Drying (Free Moisture)

- Spreading: Spread the 35 kg sample thinly on clean drying pans (coal layer loading ≤ 1 g/cm² per ISO 589).

- Drying Conditions: Allow the sample to air dry at ambient temperature (or in a fan-forced oven not exceeding 40 °C) until "visibly dry."
Note / Rationale: The drying temperature must strictly NOT exceed 40 °C to prevent coal oxidation (especially in lower-rank coals). Temperatures > 40 °C promote chemical oxygen chemisorption and mass gain, resulting in severe negative analytical bias.

- Stability Check: Weigh the pans back on LIMS. Air drying is complete when the mass loss between successive checks is less than 0.1 g (< 0.1 g).

- Final Air-Dried Mass (m_2): Record the final air-dried mass (m_2) on LIMS and AS015TF-02.

## 3. Sub-Sampling and Crushing Sequence (ISO 13909-4 / AzureSky Method)

3.1. Crushing Sequence

- Step 1 (Primary Crush): Pass the air-dried sample through the 25 mm crusher.

- Step 2 (Division): Use a clean Riffle Splitter to reduce the mass in between stage crushing.

- Step 3 (Secondary Crush):

- For Single-Stage: Crush to -10 mm.

- For Two-Stage: Crush to -3 mm.

3.2. Final Preparation

•Step 4 (Riffling): Riffle the crushed material to achieve a sample of 1 kg per riffled side.

•Step 5 (Duplicate Pans): Prepare two clean, dry moisture pans. Tare each pan on the LIMS-connected balance.

- Transfer & Initial Mass (m_3): Transfer the 1 kg sub-sample into the tared pans and immediately capture the initial net mass (m_3) on LIMS for both Pan A and Pan B.

## 4. Oven Drying (Residual and Single-Stage Moisture)

4.1. Two-Stage Residual Moisture (-3 mm AzureSky Macro Procedure)

- Place duplicate pans in the oven pre-heated to 105 °C to 110 °C (nominal setpoint ≤ 107 °C).

- Initial Interval: Dry continuously for 3 hours.

- First Weighing: Capture mass on LIMS after 3 hours, then return to oven for 30 minutes.

- Stability Criteria: If mass loss between consecutive 30-minute intervals is ≤ 0.1 g, the sample will read "stable" in LIMS, signaling that drying is complete. Capture final stabilized mass (m_4) on LIMS.

- Failure Limit: If the sample does not stabilize within 5 hours, the test is marked as "Failed" in LIMS, the test is invalidated, and it must be repeated completely using a backup split.

4.2. Single-Stage Moisture (-10 mm)

- Duplicate 1 kg portions crushed to -10 mm are placed in the oven at 105 °C to 110 °C.

- The sample remains in the oven for the full 5-hour duration continuously, without the 3-hour initial lapse check.

- Remove pans and weigh hot as soon as possible (within 5 minutes per ISO 589 Cl. 8.2.1) to avoid atmospheric moisture re-adsorption.

- Capture final stabilized mass on LIMS and AS015TF-01.

## 5. Calculations

5.1. Surface (Free) Moisture (M_ar) — Two-Stage Stage 1

mar=m1-m2m1  x 100

Where:

- m_1 = Initial mass of bulk sample as received (kg, per AS015TF-02).

- m_2 = Mass of bulk sample after Stage 1 air drying (kg, per AS015TF-02).

Calculation Example (Assessment Q10):
Given m_1 = 35.00 kg (as received) and m_2 = 32.55 kg (air dried):
M_ar = [(35.00 - 32.55) / 35.00] × 100 = (2.45 / 35.00) × 100 = 7.00%

5.2. Residual Moisture (M_res) — Two-Stage Stage 2

Mres=m3-m4m3 X 100

Where:

- m_3 = Net sample mass before oven drying (~1000 g net charge, per AS015TF-02).

- m_4 = Final net stabilized sample mass after oven drying (g, per AS015TF-02).

5.3. Total Moisture (TM)

5.3. Total Moisture (TM) Equations

A. Two-Stage Total Moisture (ISO 589 Equation 3):
TM = M_ar + [(100 - M_ar) × M_res] / 100
Where M_ar is surface moisture (%) and M_res is residual moisture (%).

B. Single-Stage Total Moisture (ISO 589:2008 Clause 8.2.3 Equation 5):
TM = [(m_2 - m_3) / (m_2 - m_1)] × 100

Where (strictly standardized per AS015TF-01 and ISO 589 Cl. 8.2.3):
• m_1 = Mass of empty, clean, dry drying tray tare (g)
• m_2 = Gross mass of drying tray plus test sample before oven drying (g)
• m_3 = Gross mass of drying tray plus test sample after oven drying (g)
Note: In Single-Stage determination, moisture loss is (m_2 - m_3) and initial test sample net mass is (m_2 - m_1). These variables are distinct from the Two-Stage bulk variables m_1 and m_2.

TM=mar+100-mar X mres100

## 6. Precision Limits (ISO 589)

6.1. ISO 589 Normative Precision Limits vs. AzureSky Decision Rules

A. ISO 589:2008 Normative Precision (Table 1):
• Second-Stage Moisture Repeatability (r): 0.30% (same operator, same apparatus, duplicate determinations).
• Total Moisture Repeatability (r): 0.50% (same laboratory, duplicate determinations).
• Total Moisture Reproducibility Critical Difference (R): 1.50% (different laboratories).

B. AzureSky Quality Decision Rule (ISO/IEC 17025 Clause 7.8.6 / AS010IW):
To ensure reporting validity, AzureSky applies an internal automated decision rule in LIMS: If the difference between duplicate Pan A and Pan B exceeds 0.30% for second-stage moisture (|M_res,A - M_res,B| > 0.30%) or 0.50% for total moisture, LIMS triggers a "REPEAT ALERT". The test result is invalidated, quarantined from release, and a complete repeat determination is initiated.

Worked Example (Assessment Q11):
Duplicate residual moisture results: Pan A M_res = 4.20%, Pan B M_res = 4.58%.
Difference: |4.58% - 4.20%| = 0.38%. Since 0.38% > 0.30%, the duplicate pair exceeds the repeatability limit. LIMS triggers a REPEAT ALERT and the sample must be re-tested.

| Property | Repeatability | Reproducibility |
| --- | --- | --- |
| Second Stage moisture | 0.3 | Not Applicable |
| Total moisture | 0.5 | 1.5 |

Note on LIMS vs. ISO 589 Formula Alignment (Assessment Q12):
LIMS calculates total moisture by averaging the duplicate residual moistures first (M_res,avg) and inserting it directly into the TM formula: TM = m_ar + [(100 - m_ar) × M_res,avg] / 100.
In contrast, the ISO 589 standard calculates total moisture for Pan A and Pan B individually and averages them at the end.
Because the equation is mathematically linear, both methods yield identical numerical results, and both methods are fully accepted at AzureSky.

## 7. Equipment Verification and Quality Control

To ensure metrological traceability and measurement validity under ISO/IEC 17025:2017 Clause 6.4 and Clause 7.7, equipment verification must strictly adhere to the approved laboratory control forms:

1. Moisture Drying Oven (Ref: AS023TF Rev 00):
• Frequency: Daily before use.
• Verification Method: Calibrated thermocouple reading (Average of 3 readings + Thermocouple Correction Factor) evaluated against the oven indicator set point and verified within the working tolerance limit of 105 °C to 110 °C.

2. Crusher Performance & Sieve Efficiencies (Ref: AS001IW Rev 04 Clause 4.1):
• Coarse Crushers (25 mm and 10 mm): Efficiency check conducted every Monday at the beginning of the shift before work commences.
• Fine Crushers (3 mm and 4.75 mm): Efficiency check conducted on the day of use, immediately before using the crusher.

Moisture Drying Oven (AS023TF): Calibrated thermocouple reading (Average + Correction Factor) checked daily against tolerance 105 °C to 110 °C (nominal indicator set point ≤ 107 °C).

Crusher Efficiency Checks (AS001IW Cl. 4.1):   - 25 mm & 10 mm crushers: Verified every Monday at the start of shift before work commences.   - 3 mm & 4.75 mm crushers: Verified on the day of use before use.

| ## Revision Table |  |  |
| --- | --- | --- |
| Rev number | Date | Reason for revision |
| 01 | 10/2026 | Technical revision and standardization: 1. Standardized variable notation: isolated Single-Stage variables (m1=tare, m2=gross initial, m3=gross dried per ISO 589 Cl. 8.2.3 and AS015TF-01) from Two-Stage variables. 2. Formally documented AzureSky in-house validated macro-pan procedure (1 kg at -3 mm) alongside ISO 589:2008 normative criteria. 3. Aligned bulk preparation with ISO 13909-4:2025 Table 1 (35 kg min mass at 50 mm top-size) and documented free water separation handling. 4. Incorporated ISO 589 constant-mass definitions (≤ 0.2% over ≥ 25% duration) and ambient re-equilibration requirements for air drying (≤ 40 °C). 5. Differentiated ISO 589 normative precision limits from AzureSky internal LIMS repeat decision rules (AS010IW). 6. Aligned equipment controls with AS023TF and AS001IW. |